## Traceability Matrix
Req ID | Test Cases | Test Script | Methods | Doc Test Step ID
SDAUIRS852 | Verify app displays the sensor wear duration after the warmup in home screen. Verify app displays the sensor warmup time during the sensor warmup in home screen. Verify transition from sensor warmup time to sensor wear duration in home screen. Verify the transition from days to hours, hours to mins, mins to sensor expired. Verify the graphical display of sensor wear duration & color code. | UIHomeScreen_T001_VerifySensorLife_PairingStatus_WarmUp | verifySensorWarmupForFirstSensor verifySensorlifefor14days verifySensorlifefor24hrs verifySensorlifefor60orlessMins verifySensorlifefor1Min verifyAlerts verifyReadytoscan clickYesOnNotActiveError clickNoOnNotActiveError | 1208
SDAUIRS855 | Pair with sensor which will expire soon. Verify app notifies the user within 1 min of sensor expiration. | UIHomeScreen_T001_VerifySensorLife_PairingStatus_WarmUp | verifySensorEnded | 1208
SDAUIRS856 | Change the Sensor activation date. Verify the sensor wear duration is displayed in 14 days. Verify the sensor wear duration is displayed in 2 days. | UIHomeScreen_T002_DisplaySensorLife | verifySensorlifefor4days verifySensorlifefor2daysto24hrs verifySensorlife5to4days verifyHomeScreen verifySensorlifefor14days | 1147
SDAUIRS857 | Change the Sensor activation date. Verify the sensor wear duration is displayed in 24 hours. Verify the sensor wear duration is displayed in 2 hours. | UIHomeScreen_T002_DisplaySensorLife | verifySensorlifefor2daysto24hrs verifySensorlifefor2hrs verifySensorlifefor2hrsto60Mins | 1147
SDAUIRS858 | Change the Sensor activation date. Verify the sensor wear duration is displayed in 60 minutes. Verify the sensor wear duration is displayed in 1 minute. | UIHomeScreen_T002_DisplaySensorLife | verifySensorlifefor2hrsto60Mins verifySensorlifefor30Mins verifySensorlife2minto1min | 1147
SDAUIRS859 | Verify remaining warmup duration is displayed when the sensor is in warmup state. | UIHomeScreen_T002_DisplaySensorLife | verifySensorWarmupinHomeScreen verifySensorWarmupForFirstSensor verifyReadytoscan editConfigurtionForFirstWarmupPeriod editConfigurtionForWarmupPeriod newSensorStartingUpMsg | 1147
SDAUIRS860 | Scan the sensor. Verify app displays the most recent scan result for 15 mins Within 15mins, scan the sensor again. Verify app displays the most result scan result for 15 mins | UIHomeScreen_T003_Verify_RealtimeBG_Value_Timestamp_TimeZone | verifyReadyToScanPage | 1148
SDAUIRS1221 | Scan the MOCK1 Sensor. Verify app displays the glucose UOM (mg/dL or mmol/L) when displaying the last scan result in Home screen | UIHomeScreen_T003_Verify_RealtimeBG_Value_Timestamp_TimeZone | verifyReadyToScanPage verifyLogbookScandata | 1148
SDAUIRS861 | Don't scan the sensor for 15 mins. Verify app displays the time of the most recent scan adjusted into the current timezone. Scan the sensor again & verify app displays the scan result. | UIHomeScreen_T003_Verify_RealtimeBG_Value_Timestamp_TimeZone | verify24hourFormatLastScan verify12hourFormatLastScan verifyTimeZone verifyReadyToScanPageNoValue verifyTimeAfterTimeZoneChangePSTtoMST verifyTimeAfterTimeZoneChangeISTtoPHT | 1148
SDAUIRS866 | Change the format to 12 hour Verify graph displays the X-axis for the 24-hour period ending at the current time Change the format to 24 hour Verify graph displays the X-axis for the 24-hour period ending at the current time | UIHomeScreen_T004_HomeScreen_Verify_Glucose_Graph | Manual Image Verification | 1149 1211 1212 1213 1214
SDAUIRS867* | Verify graph displays the Y-axis from 0 to at least 350 mg/dL. (0-21 mmol/L) Verify graph displays the Y-axis from 0 to 500 mg/dL if glucose data > 350 mg/dL (0- 27 mmol/L) | UIHomeScreen_T004_HomeScreen_Verify_Glucose_Graph | Manual Image Verification | 1149 1211 1212 1213 1214
SDAUIRS871* | Add a multiple notes in scan result. Verify graph displays the badge indicating the number of notes being represented by that icon. Verify the number of notes is displayed correctly. Verify the note icon time plotted in graph matches approximately with the note time. | UIHomeScreen_T004_HomeScreen_Verify_Glucose_Graph | verifyAddedNotesTime verifyAddedNotesIcon verifyAddedNotesBadge | 1149 1249
SDAUIRS1125 | Add a multiple notes in scan result. Verify graph displays the badge indicating the number of notes being represented by that icon. Verify the number of notes is displayed correctly. Verify the note icon time plotted in graph matches approximately with the note time. | UIHomeScreen_T004_HomeScreen_Verify_Glucose_Graph | verifyAddedNotesIcon verifyAddedNotesBadge | 1149
SDAUIRS872 | Verify home screen graph distinguishes the target range. Also verify target range is plotted as per the current settings. | UIHomeScreen_T005_Verify_Glucose_Target_Range | Manual Image Verification | 1210 1215 1216 1217 1218 1219 1220
SDAUIRS874 | Verify app displays the message instructing the user to scan their sensor in home screen if there is no glucose data within the past 24 hours and a sensor is currently paired | UIHomeScreen_T006_Apply_New_Sensor | verifysensorTitlesstatus verifysensorstatusDesc | 1150
SDAUIRS875 | Verify app displays the message instructing the user to scan a new sensor to activate it in home screen if there is no glucose data within the past 24 hours and no sensor is currently paired. | UIHomeScreen_T006_Apply_New_Sensor | verifyApplynewsensorscreen verifyAlerts clickYesOnNotActiveError | 1150
SDAUIRS1174 | Set the SAS Mode to Default Launch the application. Verify the Sensor pairing status displayed as 'Apply New Sensor' in Home Screen. Verify the Sensor pairing status displayed as 'Ready to Scan” in Home Screen. Add the Realtime BG value as 100 mg/dL and scan the sensor by tapping the “Scan Mock Sensor” button. Set the Sensor Activation date to less than 13 days 23 hrs 59 min prior to the current time, wait for 1 min, once the sensor life ends, Verify the Sensor pairing status displayed as 'Apply and Scan New Sensor' in Home Screen. | UIHomeScreen_T007_Sensor_pairing_status_HomeScreen_TIME_IN_TARGET_Historical_BG_AVERAGE_GlucoseGraph | verifyApplynewsensorscreen verifyReadytoscan verifySensorEnded verifySensorlifefor1Min | 1151
SDAUIRS863 | Scan the sensor with multiple in range /out range glucose values. Verify app displays the percentage of historical glucose points whose values are within (inclusive) the target glucose range, for the set of historical glucose points with UTC timestamps within the past 24 hours. Verify the displayed value correctly. Scan the sensor with no in range glucose values. Verify app displays the percentage of historical glucose points whose values are within (inclusive) the target glucose range, for the set of historical glucose points with UTC timestamps within the past 24 hours as 0%. Scan the sensor with no historic for 24 hours. Verify app displays TIME IN TARGET as "--". | UIHomeScreen_T007_Sensor_pairing_status_HomeScreen_TIME_IN_TARGET_Historical_BG_AVERAGE_GlucoseGraph | verifyHistorictimeintarget verifyTimeZone | 1151
SDAUIRS864 | Scan the sensor with glucose values. Verify app displays the average of all historical glucose points with UTC timestamps within the past 24 hours. Verify the displayed value correctly in correct UOM.. " Scan the sensor with no historic for 24 hours. Verify app displays AVERAGE as "--". | UIHomeScreen_T007_Sensor_pairing_status_HomeScreen_TIME_IN_TARGET_Historical_BG_AVERAGE_GlucoseGraph | verifyAvgHistGlucose verifyTimeZone | 1151
SDAUIRS865 | Scan the sensor with glucose values. Verify app displays the glucose graph If there is glucose data with UTC timestamps within the past 24 hours. Verify the displayed information in glucose graph. Scan the sensor with no glucose values within 24 hours. Verify app doesn't display the glucose graph. | UIHomeScreen_T007_Sensor_pairing_status_HomeScreen_TIME_IN_TARGET_Historical_BG_AVERAGE_GlucoseGraph | verifysensorTitlesstatus verifyCurrentBGValue Manual Image Verification | 1151 1247 1250 1251
SDAUIRS868 | Add a multiple historic glucose values. Scan the sensor. Verify historic glucose values are plotted correctly in graph. | UIHomeScreen_T008_GlucoseGraph_Realtime_Historical_BG_TimeZoneChange_Events_Details | Manual Image Verification | 1205 1221 1222
SDAUIRS1241 | Click the time change icon & verify timezone change information is displayed correctly. | UIHomeScreen_T008_GlucoseGraph_Realtime_Historical_BG_TimeZoneChange_Events_Details | timeChangeIndicatorInfo | 1205 1253
SDAUIRS870 | Add some historical values. Scan the sensor. Change the current timezone to different timezone. Add some historical values. Scan the sensor. Verify the timechange icon is plotted correctly in the current timezone. | UIHomeScreen_T008_GlucoseGraph_Realtime_Historical_BG_TimeZoneChange_Events_Details | verifyTimeChangeIndicator Manual Image Verification | 1205 1252
SDAUIRS869 | Scan the sensor. Verify app displays the real time glucose point in the graph & also verify real time glucose point is plotted correctly. Wait for 15 mins. Verify the real time glucose point disappears in the graph after 15 mins. | UIHomeScreen_T008_GlucoseGraph_Realtime_Historical_BG_TimeZoneChange_Events_Details | Manual Image Verification | 1205 1223 1224 1226 1225
SDAUIRS873 | Add two or three historic values within 20 mins gap. Verify historic glucose values are connected. Add two or three historic values > 20 mins gap. Verify historic glucose values are not connected. Add two or three historic values exactly at 20 mins gap.Verify historic glucose values are connected. | UIHomeScreen_T009_GlucoseGraph_HistoricBGScans_In20Mins | Manual Image Verification | 1176 1127 1128 1129
SDAUIRS1232 | Make sure historic glucose values time gap & scan result time gap are within 35 mins. Verify app connects historic glucose values & scan result in glucose graph. Make sure historic glucose values time gap & scan result time gap are exactly 35 mins. Verify app connect historic glucose values & scan result in glucose graph. Make sure historic glucose values time gap & scan result time gap are > 35 mins. Verify app doesn't connect historic glucose values & scan result in glucose graph. | UIHomeScreen_T010_GlucoseGraph_Realtime_Historical_BG_PointsConnection | Manual Image Verification | 1177 1230 1239 1242
SDAUIRS1230 | "Make sure historic glucose values time gap & scan result time gap are within 35 mins. Verify app connects historic glucose values & scan result in home screen glucose graph " "Make sure historic glucose values time gap & scan result time gap are exactly 35 mins. Verify app connect historic glucose values & scan result in home screen glucose graph " "Make sure historic glucose values time gap & scan result time gap are > 35 mins. Verify app doesn't connect historic glucose values & scan result in home screen glucose graph. | UIHomeScreen_T010_GlucoseGraph_Realtime_Historical_BG_PointsConnection | Manual Image Verification | 1177 1231 1240 1243
SDAUIRS1231 | Make sure historic glucose values time gap & scan result time gap are within 35 mins. Verify app connects historic glucose values & scan result in Logbook glucose graph Make sure historic glucose values time gap & scan result time gap are exactly 35 mins. Verify app connect historic glucose values & scan result in Logbook glucose graph. Make sure historic glucose values time gap & scan result time gap are > 35 mins. Verify app doesn't connects historic glucose values & scan result in Logbook glucose graph. | UIHomeScreen_T010_GlucoseGraph_Realtime_Historical_BG_PointsConnection | Manual Image Verification | 1177 1232 1241 1244

## Methods - NOT Req Specific
Test Script Name | Methods - NOT Req Specific
UIHomeScreen_T001_VerifySensorLife_PairingStatus_WarmUp | debugDrawerClearData selectingSASMode editConfigurtionForFirstWarmupPeriod closeDebugDrawer advanceTime launchAppWithdata navigateToLogBook navigateToHomePage setToDefault currentSystemTime debugDrawerClearData clickBackArrowORMenu addRealtimeMockSensor scanMockSensor changePhoneHourTimeFormat
UIHomeScreen_T002_DisplaySensorLife | setToDefault openDebugDrawer debugDrawerClearData selectingSASMode bgValueWithStaticTime scanMockSensor
UIHomeScreen_T003_Verify_RealtimeBG_Value_Timestamp_TimeZone | navigatetoHomeScreen changePhoneHourTimeFormat appNetworkTimeCheck changeTimeZone debugDrawerClearData selectingSASMode openDebugDrawer bgValueWithStaticTime scanMockSensor navigateToHomeScreenFromMyGlucosePage closeDebugDrawer navigateToLogBook launchAppWithdata timezoneCleanUp
UIHomeScreen_T004_HomeScreen_Verify_Glucose_Graph | clickAddNote navigateToLogBook changePhoneHourTimeFormat setToDefault debugDrawerClearData selectingSASMode openDebugDrawer bgValueWithStaticTime scanMockSensor navigateToHomeScreenFromMyGlucosePage closeDebugDrawer navigatetoHomeScreen setAdvancedLogBookTime addNotetoFoodOptions addNotetoRAIoptions addNotetoLAIoptions addNoteExercise addNoteDoneButton cleanUpT004 launchAppWithdata
UIHomeScreen_T005_Verify_Glucose_Target_Range | changePhoneHourTimeFormat debugDrawerClearData selectingSASMode openDebugDrawer bgValueWithStaticTime navigateToHomeScreenFromMyGlucosePage scanMockSensor closeDebugDrawer setTargetGlucoseRange setToDefault
UIHomeScreen_T006_Apply_New_Sensor | createAccount changePhoneHourTimeFormat selectingSASMode editConfigurtionForWarmupPeriod openDebugDrawer bgValueWithStaticTime scanNowWithOutScanResults navigateToHomeScreenFromMyGlucosePage startUp verifySensorlifefor1Min advanceTime selectingSASMode cleanUpT006 debugDrawerClearData verifySensorlife currentSystemTime
UIHomeScreen_T007_Sensor_pairing_status_HomeScreen_TIME_IN_TARGET_Historical_BG_AVERAGE_GlucoseGraph | createAccount selectingSASMode openDebugDrawer bgValueWithStaticTime scanMockSensor navigateToHomeScreenFromMyGlucosePage closeDebugDrawer appNetworkTimeCheck debugDrawerClearData setTargetGlucoseRange verifylast24Hour advanceDeviceTime navigateToLogBook navigateToHomePage changeTimeZone verifyTimeAfterTimeZoneChangePSTtoCST verifyTimeAfterTimeZoneChangeISTtoPHT launchAppWithdata setTheTime setToDefault cleanUpT007 startUp timezoneCleanUp
UIHomeScreen_T008_GlucoseGraph_Realtime_Historical_BG_TimeZoneChange_Events_Details | changePhoneHourTimeFormat appNetworkTimeCheck changeTimeZone currentSystemTime debugDrawerClearData setToDefault closeDebugDrawer verifyReadytoscan verifyNoHistoricalData openDebugDrawer navigateToHomeScreenFromMyGlucosePage scanNowWithOutScanResults bgvalueWithStaticTime launchAppWithdata advanceDeviceTime navigateToLogBook navigateToHomePage selectingSASMode timezoneCleanUp
UIHomeScreen_T009_GlucoseGraph_HistoricBGScans | debugDrawerClearData cleanUpT009 setTargetGlucoseRange changePhoneHourTimeFormat setToDefault selectingSASMode openDebugDrawer bgValueWithStaticTime scanMockSensor navigateToHomeScreenFromMyGlucosePage closeDebugDrawer defineHistoricValue
UIHomeScreen_T010_GlucoseGraph_Realtime_Historical_BG_PointsConnection | setToDefault selectingSASMode appNetworkTimeCheck navigateToAbout getSmartDeviceDetails clickBackArrowORMenu setTheTime changePhoneHourTimeFormat debugDrawerClearData openDebugDrawer defineHistoricValue defineRealtimeValue scanMockSensor navigateToHomeScreenFromMyGlucosePage closeDebugDrawer navigateToLogBook verifyLogbookDetails selectLogbookValue