## Traceability Matrix
Req ID | Test Cases | Test Script Name | Script Method Name | DVT Protocol Test Step ID
SDASRS257 | The software shall log a code whenever Real Time sensor glucose cannot be calculated according to the algorithm. The software shall provide a means to review the error and event log. Load the dataset. Verify the following error & events are listed in the event log. NOT_ACTIVE TEMP_HIGH TEMP_LOW TEMPORARY_PROBLEM NOT_COMPATIBLE IN_WARMUP RESPONSE_CORRUPT | SensorScanning_T001_Mock1UI_Execptions_without_NFCApp_installed SensorScanning_T002_Mock1UI_TransmissionError_with_NFCApp_installed SensorScanning_T003_Mock1UI_Exceptions_Change_Mock_Sensor_State | clickOnEventLog verifyEventLogContent eventlogVerifyErrIcon | 1205 1176 1177
SDASRS330 | Scan the Mock sensor & verify app displays the sensor too hot message to the user. Scan the Mock sensor & verify app displays the sensor too cold message to the user. | SensorScanning_T001_Mock1UI_Execptions_without_NFCApp_installed | scanAlerts verifyTEMPHIGHSensorScreen verifyTEMPLOWSensorScreen clickOKSensorReady | 1205
SDASRS368 | The software shall log a code whenever Real Time sensor glucose cannot be calculated according to the algorithm. The software shall provide a means to review the error and event log." Load the dataset. Verify the following error & events are listed in the event log. NOT_ACTIVE TEMP_HIGH TEMP_LOW TEMPORARY_PROBLEM NOT_COMPATIBLE IN_WARMUP RESPONSE_CORRUPT | SensorScanning_T001_Mock1UI_Execptions_without_NFCApp_installed | verifyEventLogContent | 1205
SDASRS827 * | Make sure sensor is in Temporary problem state. Scan the sensor. Verify app notifies the user of the failure and discontinue further calculation, storage, transmission or display of these data | SensorScanning_T001_Mock1UI_Execptions_without_NFCApp_installed | verifyTEMPORARYPROBLEMSensorScreen | 1205
SDAUIRS1083 | While scanning is in progress move the sensor out of range. Verify app displays an error message & sensor is not activated successfully. | SensorActivation_T002_Mock1UI_TranmissionError_with_NFCApp_installed SensorActivation_T001_Mock1UI_Exceptions_without_NFCApp_installed | scanAlerts verifySensorScreenWithOtherNFCApp verifyTransmissionSensorScreen clickSensorOk verifyHomeScreenTitle | 1150 1151
SDAUIRS1084 | Bring expired sensor in range with Phone. Verify app displays an error message | SensorActivation_T001_Mock1UI_Exceptions_without_NFCApp_installed | scanAlerts verifyExpiredSensorScreen clickSensorOk homePage | 1150
SDAUIRS1085 | Bring sensor in range with Phone1. Verify sensor is paired successfully with Phone1. Now bring the same sensor in range with Phone2. Verify app displays an error message to the user & Phone2 is not paired with sensor. | SensorActivation_T001_Mock1UI_Exceptions_without_NFCApp_installed | scanAlerts clickNotActive verifyStartedSensorScreen clickSensorOk verifyHomeScreenTitle | 1150
SDAUIRS1087 | Make sure sensor is with response corrupt data. Scan the sensor & verify app displays the error message & sensor is not paired with Phone. | SensorActivation_T001_Mock1UI_Exceptions_without_NFCApp_installed | scanAlerts verifyResponseCorruptSensorScreen clickSensorOk verifyHomeScreenTitle | 1150
SDAUIRS1088 | Make sure sensor is in error terminated state. Scan the sensor & verify app displays the error message & sensor is not paired with Phone. | SensorActivation_T001_Mock1UI_Exceptions_without_NFCApp_installed | scanAlerts verifyTerminatedSensorScreen clickSensorOk verifyHomeScreenTitle | 1150
SDAUIRS1089 | Make sure sensor is in Sensor removed state. Scan the sensor & verify app displays the error message to the user & sensor is not paired with Phone. | SensorActivation_T001_Mock1UI_Exceptions_without_NFCApp_installed | scanAlerts verifyTerminatedSensorScreen clickSensorOk verifyHomeScreenTitle | 1150
SDAUIRS1090 | Bring the sensor(other than librelink supported sensor) in range with Phone. Scan the sensor & verify app displays the error message to the user & sensor is not paired with Phone. | SensorActivation_T001_Mock1UI_Exceptions_without_NFCApp_installed | scanAlerts verifyNotCompatableSensorScreen clickSensorOk verifyHomeScreenTitle | 1150
SDAUIRS1096 | Bring sensor 1 in range with Phone & scan the sensor. Verify app displays the glucose result. Bring sensor2(not activated before) in range with same Phone. Verify app prompts the user whether to switch sensors. Verify result screen is not displayed. | SensorScanning_T001_Mock1UI_Execptions_without_NFCApp_installed | scanAlerts verifyNotActiveSensorScreen clickNotActive clickBackToHomeScreen | 1205
SDAUIRS1097 | Move the sensor out of range during scanning in progress & verify app displays the scan failure message. Verify result screen is not displayed. | SensorScanning_T002_Mock1UI_TransmissionError_with_NFCApp_installed | scanAlerts verifySensorScreenWithOtherNFCApp clickOKSensorReady | 1176
SDAUIRS1098 | Make sure sensor is in Error terminated state. Scan the sensor & verify app displays the error message. Verify result screen is not displayed. | SensorScanning_T003_Mock1UI_Exceptions_Change_Mock_Sensor_State | scanAlerts verifyTerminatedSensorScreen clickOkErrorDialog | 1177
SDAUIRS1099 | Make sure sensor with Response Corrupt data. Scan the sensor & verify app displays the error message. Verify result screen is not displayed. | SensorScanning_T001_Mock1UI_Execptions_without_NFCApp_installed | scanAlerts verifyRESPONSECORRUPTSensorScreen clickOKSensorReady | 1205
SDAUIRS1101 | Make sure sensor 1 is in warmup state & paired with Phone1. Bring new sensor2 & pair with Phone1. Now bring sensor 1 in range with Phone 1. Verify app displays the error message to the user. Verify result screen is not displayed. | SensorScanning_T001_Mock1UI_Execptions_without_NFCApp_installed | scanAlerts verifyINWARMUPSensorScreen clickOKSensorReady | 1205
SDAUIRS1102 | Make sure sensor is expired. Scan the expired sensor & verify app displays the error message to the user. Verify result screen is not displayed. | SensorScanning_T003_Mock1UI_Exceptions_Change_Mock_Sensor_State | scanAlerts verifyExpiredSensorScreen clickOkErrorDialog | 1177
SDAUIRS1103 | Make sure sensor temperature is too high. Scan the sensor & verify app displays the error message to the user. Verify result screen is not displayed. | SensorScanning_T001_Mock1UI_Execptions_without_NFCApp_installed | scanAlerts verifyTEMPHIGHSensorScreen clickOKSensorReady | 1205
SDAUIRS1104 | Make sure sensor temperature is too low. Scan the sensor & verify app displays the error message to the user. Verify result screen is not displayed. | SensorScanning_T001_Mock1UI_Execptions_without_NFCApp_installed | scanAlerts verifyTEMPLOWSensorScreen clickOKSensorReady | 1205
SDAUIRS1105 | Make sure sensor has the data quality problem. Scan the sensor & verify app displays the error message to the user. Verify result screen is not displayed. | SensorScanning_T001_Mock1UI_Execptions_without_NFCApp_installed | scanAlerts verifyTEMPORARYPROBLEMSensorScreen clickOKSensorReady | 1205
SDAUIRS1106 | Make sure sensor is in Insertion failure state. Scan the sensor & verify app displays the error message to the user. Verify result screen is not displayed. | SensorScanning_T003_Mock1UI_Exceptions_Change_Mock_Sensor_State | scanAlerts verifyINSERTIONFAILURESensorScreen | 1177
SDAUIRS1107 | Make sure sensor is in sensor removal state. Scan the sensor & verify app displays the error message to the user. Verify result screen is not displayed. | SensorScanning_T003_Mock1UI_Exceptions_Change_Mock_Sensor_State | scanAlerts verifyTerminatedSensorScreen clickOkErrorDialog | 1177
SDAUIRS1108 | Bring the Libre Pro sensor in range with Phone. Scan the sensor & verify app displays the error message to the user. Verify result screen is not displayed. | SensorScanning_T001_Mock1UI_Execptions_without_NFCApp_installed | scanAlerts verifyNotCompatableSensorScreen clickOKSensorReady | 1205
SDAUIRS1174 | Set the SAS Mode to Default Launch the application. Verify the Sensor pairing status displayed as 'Apply New Sensor' in Home Screen. Verify the Sensor pairing status displayed as 'Ready to Scan” in Home Screen. Add the Realtime BG value as 100 mg/dL and scan the sensor by tapping the “Scan Mock Sensor” button. Set the Sensor Activation date to less than 13 days 23 hrs 59 min prior to the current time, wait for 1 min, once the sensor life ends, Verify the Sensor pairing status displayed as 'Apply and Scan New Sensor' in Home Screen. | SensorActivation_T002_Mock1UI_TranmissionError_with_NFCApp_installed SensorActivation_T001_Mock1UI_Exceptions_without_NFCApp_installed | editConfigurationTime verifyHomeScreenTitle | 1150 1151
SDAUIRS1144 SDAUIRS1253 | The software shall provide a means for the user to choose a target glucose range such that the upper target is at least 1 mg/dL greater than the lower target when configured for mg/dL. The software shall provide a means for the user to choose a target glucose range such that the upper target is at least 0.1 mmol/L greater than the lower target when configured for mmol/L. | Settings_T001_AppSettings | clickTargetGlucoseRange navigateToAppSetting verifyUpperGlucoseRangeUnits verifyLowerGlucoseRangeUnits verifyAppSetings verifyTargetRange setGlucoseRange | 1148
SDAUIRS1145 SDASRS830 | The software shall provide a means for the user to choose grams or servings as their carbohydrate unit of measure. The software shall provide a means to change carbohydrate unit of measure settings. | Settings_T001_AppSettings | NavigateCarbohydrateUnits clickAndVerifyGramsPage clickAndVerifyServingsPage verifyAppSetings | 1148
SDAUIRS1146 | When the user has chosen servings as their carbohydrate unit of measure, the software shall provide a means to allow the user to specify a serving size between 10 and 15 grams, in 0.5 gram increments. | Settings_T001_AppSettings | NavigateCarbohydrateUnits clickAndVerifyGramsPage clickAndVerifyServingsPage | 1148
SDASRS172 | The software shall provide a means to change the Target Glucose Range. | Settings_T001_AppSettings | verifyUpperGlucoseRangeUnits verifyLowerGlucoseRangeUnits clickTargetGlucoseRange verifyTargetRange setGlucoseRange | 1148
SDAUIRS1245 | The software shall provide a means to enable text-to-speech. | Settings_T001_AppSettings | navigateToTextToSpeechPage selectspeechOption clickSave verifySpeechSettings verifyAppSetings | 1148
SDAUIRS1246 | The software shall provide a means to disable text-to-speech. | Settings_T001_AppSettings | navigateToTextToSpeechPage selectspeechOption clickSave verifySpeechSettings verifyAppSetings | 1148
SDAUIRS1216 | The software shall provide a means for the user to choose whether a sound should be played while scanning a sensor, if supported by the OS. Note: only applies to Android 5.0 and later | Settings_T001_AppSettings | chooseSoundorVibration clickSave verifySoundSettings navigateToSoundSettingsPage verifyScanSoundsPageDetails navigateToSoundSettingsPage verifyAppSetings | 1148
SDASRS1178 | The software shall provide a means to change scan sounds settings. | Settings_T001_AppSettings | chooseSoundorVibration clickSave verifySoundSettings navigateToSoundSettingsPage | 1148
SDAUIRS925 | The software shall provide a means to change connected account settings. | Settings_T002_MinorAccountSettings | verifyLibreLinkLaunch navigateAccountSettings verifyMinorAccountSettings clickDOB enterAgePickDOB sendValidPasswordAndSubmit verifyMinorMessage clickMinorMessageOk verifyApplicationAcceptDate verifyAccountPageEditable sendValidPasswordAndSubmit clickMinorMessageCancel verifyMinorAccountSettings verifyMinorWarningMessage | 1149
SDAUIRS926 | The software shall provide a means to change the connected account password. | Settings_T002_MinorAccountSettings | navigateToAccountPassword verifyAccountPasswordPage changePassword accountPasswordErrors homePage signInClick verifyUOMPage | 1149
SDAUIRS1238 | The first-time startup wizard shall provide a means to display a connected account’s registration pages and require successful account creation before continuing. | Settings_T002_MinorAccountSettings | createMinorAccountDetails clickNextOnAccountPage verifyMinorMessage clickMinorMessageOk verifyParentGuardianAccountpage createAccountNames createParentLogin verifyUOMPage verifyMinorWarningMessage | 1149
SDAUIRS925 | The software shall provide a means to change connected account settings. | Settings_T003_AdultAccountSettings | clickDOB navigateAccountSettings verifyAccountSettings enterAgePickDOB verifyAccountError clickAccountErrorOK sendValidPasswordAndSubmit sendValidNameandEmail sendValidName | 1211
SDAUIRS926 | The software shall provide a means to change the connected account password. | Settings_T003_AdultAccountSettings | navigateToAccountPassword verifyAccountPasswordPage changePassword homePage accountPasswordErrors | 1211
SDAUIRS1255 | The software shall validate new passwords which must be 8-36 characters and contain at least 1 digit,1 upper case letter, and 1 lower case letter. | Settings_T003_AdultAccountSettings | changePassword sendValidPasswordAndSubmit verifyAccountError accountPasswordErrors | 1211
SDAUIRS912 | The software shall provide a means to create a Reminder. | Reminder_T001_Alarm_Reminder_UI_Check | clickAddReminderandSetAlarm verifyReminder | 1208
SDAUIRS1156 | The software shall provide a means to view Reminders. | Reminder_T001_Alarm_Reminder_UI_Check Reminder_T001_Alarm_Reminder_UI_Check_2 | navigateToPage clickReminderIcon verifyReminder clickAddReminderandSetAlarm | 1208 1214
SDAUIRS260 | The software shall provide a means to edit a Reminder. | Reminder_T001_Alarm_Reminder_UI_Check | editReminder | 1208
SDAUIRS1123 | The software shall provide a means to delete a Reminder. | Reminder_T001_Alarm_Reminder_UI_Check | deleteReminder | 1208
SDAUIRS913 | The software shall provide a means to store at least 12 reminders. | Reminder_T001_Alarm_Reminder_UI_Check | clickAddReminderandSetAlarm | 1208
SDAUIRS914 | The software shall provide a means to set the time of a reminder. | Reminder_T001_Alarm_Reminder_UI_Check | clickReminderIcon clickAddReminderandSetAlarm | 1208
SDAUIRS1203 | The software shall display a message indicating which reminder is next. | Reminder_T001_Alarm_Reminder_UI_Check | verifyNextReminderwithBackGroundColor | 1208
SDAUIRS262 | The software shall display a Reminder message when that reminder becomes current. | Reminder_T001_Alarm_Reminder_UI_Check | verifyAndClickTriggerNotification verifyLibreLinkLaunch verifyEnableReminderTrigger editReminderWithTime verifyEnableReminder verifyWaitReminderTriggerNotification | 1208
SDAUIRS915 | The software shall provide a means to set the message for a reminder of up to 50 characters. | Reminder_T001_Alarm_Reminder_UI_Check | selectReminderWithFiftyPlusCharacters reminderWithFiftyPlusCharacters deleteReminderWithFiftyPlusCharacters | 1208
SDAUIRS1155 | The software shall disallow entry of newline characters in reminder messages. | Reminder_T001_Alarm_Reminder_UI_Check | reminderWithFiftyPlusCharacters | 1208
SDAUIRS916 | The software shall provide a means to set the repeat mode of a reminder to repeat at a selected interval or only once. | Reminder_T001_Alarm_Reminder_UI_Check_2 | verifyReminder enableRepeatingOptions verifyRepeatingOptions disableRepeatingOptions enableRepeatingOptions checkOneday selectAllCheckBox selectAllday verifyWaitReminderRecursive | 1214
SDAUIRS917 | The software shall provide a means to set a reminder to repeat on one or more days of the week. | Reminder_T001_Alarm_Reminder_UI_Check_2 | thisMonday thisTuesday thisWednesday thisThrusday thisFriday thisSaturday thisSunday selectMondayAfterOneWeek editReminderWithOneDay | 1214
SDAUIRS918 | The software shall provide a means to enable a reminder. | Reminder_T001_Alarm_Reminder_UI_Check | enableTrigger | 1208
SDAUIRS1183 | The software shall provide a means to disable a reminder. | Reminder_T001_Alarm_Reminder_UI_Check | verifyDisableReminderTrigger disableTrigger disableRepeatingOptions | 1208
SDAUIRS1158 | The software shall provide a means to dismiss the Reminder. | Reminder_T001_Alarm_Reminder_UI_Check_2 | clearReminderTrigger | 1214
SDAUIRS920 | The software shall provide a means to enable a timer reminder. | Reminder_T002_Timer_Reminder_UI_Check | enableScanSensor verifyEnabledToggleButton displaySensorTimer disableScanSensor clickReminderIcon verifyRemindersScanSensor | 1147
SDAUIRS1184 | The software shall provide a means to disable a timer reminder. | Reminder_T002_Timer_Reminder_UI_Check | disableScanSensor displaySensorTimer | 1147
SDAUIRS922 | The software shall provide a means for the user to set the timer reminder countdown interval in hourly increments from 1 to 12 hours. | Reminder_T002_Timer_Reminder_UI_Check | clickScanSensorandVerifyHours scanSensorTimerReminder displaySensorTimer enableScanSensor scanMockWithBGTypeAndVerifyDeaultReminder disableScanSensor setHourAndscanSensorTimerReminder | 1147
SDAUIRS923 | The default timer reminder shall remind the user to scan their sensor. | Reminder_T002_Timer_Reminder_UI_Check | verifyTimerExpired scanMockWithBGTypeAndVerifyDeaultReminder displaySensorTimer setHourAndscanSensorTimerReminder | 1147
SDAUIRS1256 | The default timer reminder shall reset after a successful scan | Reminder_T002_Timer_Reminder_UI_Check | scanMockWithBGTypeAndVerifyDeaultReminder | 1147
SDASRS1156 | The software shall provide for user management of reminders. | Reminder_T002_Timer_Reminder_UI_Check Reminder_T001_Alarm_Reminder_UI_Check | navigateToPage verifyRemindersPage verifyReminder verifyNextReminderwithBackGroundColor clickReminderIcon disableScanSensor setHourAndscanSensorTimerReminder verifyRemindersScanSensor | 1147 1208

## Methods - NOT Req Specific
Test Script Name | Methods - NOT Req Specific
SensorActivation_T001_Mock1UI_Exceptions_without_NFCApp_installed | openDebugDrawer addMockSensorData selectingSASMode debugDrawerClearData setToDefault
SensorActivation_T002_Mock1UI_TranmissionError_with_NFCApp_installed | selectingSASMode debugDrawerClearData installUninstallOtherNFCAPP launch startUp
SensorScanning_T001_Mock1UI_Execptions_without_NFCApp_installed | eventlogScrollUp startUp selectingSASMode scanMockSensorandGoBack verifyReadyToScanPage addMockSensorData debugDrawerClearData setTheTime changePhoneHourTimeFormat openDebugDrawer navigateToHelpPage
SensorScanning_T002_Mock1UI_TransmissionError_with_NFCApp_installed | currentSystemTime selectingSASMode debugDrawerClearData openDebugDrawer installUninstallOtherNFCAPP launch startUp changePhoneHourTimeFormat setTheTime addMockSensorData scanMockSensorandGoBack verifyReadyToScanPage navigateToHelpPage
SensorScanning_T003_Mock1UI_Exceptions_Change_Mock_Sensor_State | eventlogScrollUp startUp selectingSASMode scanMockSensorandGoBack verifyReadyToScanPage addMockSensorData debugDrawerClearData setTheTime changePhoneHourTimeFormat openDebugDrawer navigateToHelpPage
Settings_T001_AppSettings | cleanUpT001 debugDrawerClearData verifyTargetGlucoserange clickBackArrowORMenu setCarboUnits clickBackArrowORMenu
Settings_T002_MinorAccountSettings | clearData closeDebugDrawer clickOnNext getStarted clickBackArrowORMenu termsOfPageAccept clickAndAcceptPrivacyNoticePage defaultSettings unInstall install launch
Settings_T003_AdultAccountSettings | serverErrorFound
Reminder_T001_Alarm_Reminder_UI_Check Reminder_T001_Alarm_Reminder_UI_Check_2 | deleteReminderIfAvailable appNetworkTimeCheck startUp openSetting reminderAdvancedtime navigateToHomePage clearAlarmExpiredNotification vanish closeNotification currentSystemTime debugDrawerClearData changePhoneHourTimeFormat reLaunch backToReminder setSystemTime advanceStaticTime setReminderName CleanUpT001
Reminder_T002_Timer_Reminder_UI_Check | navigateToPage startUp navigateToHomePage appNetworkTimeCheck selectingSASMode changePhoneHourTimeFormat closeSensorTimeElements clickBackArrowORMenu closeRunningApplication launchAppWithdata backToReminder currentSystemTime CleanUpT002