## Traceability Matrix
Req ID | Test Cases | Test Script Name | Script Method Name | DVT Protocol Test Step ID
SDAUIRS877 | Verify app displays the LogBook in the Menu list. Click on LogBook in the Menu list & verify Logbook screen is displayed. | UILogbook_T001_Logbook_Navigation | navigateToLogBook verifyLogBookMenu verifyLogbookPageDisplayed verifyaddedNote | 1208
SDASRS1162 | Verify app displays the LogBook in the Menu list. The software shall provide for user management of a logbook. | UILogbook_T001_Logbook_Navigation | navigateToLogBook verifyLogBookMenu verifyLogBookDisplaysCurrentDate verifyLogBookEmpty verifyLogbookPageContents verifyBloodDropIcon verifyaddedNote | 1208
SDAUIRS883 | Set the Phone time format to 12 hours. Verify X-axis is displayed from midnight to midnight. Set the Phone time format to 24 hours. Verify X-axis is displayed from midnight to midnight. | UILogbook_T002_Logbook_Graph_Axis | navigateToLogBook selectNotes verifyLogbookDetail verifyReportXAxisWith12HoursFormat verifyReportXAxisWith24HoursFormat | 1147
SDAUIRS884 | Verify graph displays the Y-axis from 0 to at least 350 mg/dL. (0-21 mmol/L) Verify graph displays the Y-axis from 0 to 500 mg/dL if glucose data > 350 mg/dL (0- 27 mmol/L) | UILogbook_T002_Logbook_Graph_Axis | navigateToLogBook selectNotes verifyLogbookDetail Manual Image Verification | 1147 1215 1216 1217
SDAUIRS878 | Scan the sensor & verify app displays the scan result in Log book Screen. Click on scan results in Log book & verify scan results information are displayed correctly in Log book Detail Page. | UILogbook_T004_Logbook_Scan_Note_Verification | selectingSASMode addMockSensorDatawithStaticTimeAndDate navigateToLogBook verifyLastScan verifyLogBookDisplaysCurrentDate verifyLogbookNoteDetails verifyScantime verifyLogbookScanDetail verifyCommentLogbookDetailPage | 1148
SDAUIRS880 | Change the Phone time zone between one scan in PST & one scan in EST . Verify app provides the information about time changes in Log book screen. | UILogbook_T005_Logbook_Change_Time_Zone | addRealtimeMockSensor addMockSensorDatawithStaticTimeAndDate scanMockSensor verifyLastScan navigateToLogBook verifyLogbookNotes verifyLogbookInformativeTextPSTtoEST selectLogbookNote verifyLogbookDetailsInformativeTextPSTtoEST verifyLogbookInformativeTextESTtoPST verifyLogbookDetailsInformativeTextESTtoPST Manual Image Verification | 1149
SDAUIRS889 | Add a historic data. Scan the sensor. Change the current timezone to different timezone. Add a historic data. Scan the sensor. Verify timechange icon is displayed. Also verify the timechange icon is plotted correctly. | UILogbook_T006_Logbook_Graph_Time_Zone_changes | addRealtimeMockSensor bgValueWithStaticTime scanMockSensor navigateToLogBook selectNotes verifyTimeChangeIcon verifyLogbookDetail Manual Image Verification | 1150 1220 1219
SDAUIRS1242 | Click the time change icon & verify timezone change information is displayed correctly. | UILogbook_T006_Logbook_Graph_Time_Zone_changes | clickOnTimeChangeIcon verifyTimezoneChangeBetweenLast2GlucoseReadings Manual Image Verification | 1150 1220 1219
SDAUIRS882 | Scan the sensor. Click the Calendar & select each day & verify app displays the list of scan results, time change events and notes for each day( if there is data collected for that day) in the period beginning 89 days before the current day and ending on the current day in Log book screen. Verify the information is displayed correctly in Logbook screen. Click on the Note & scan results and verify app displays the glucose graph in Logbook detail screen. Verify the information is displayed correctly. Verify If there is no real time data on that day verify app displays 'No entries exist for this date' | UILogbook_T007_Logbook_Glucose_Graph_for_90_days_Auto | navigateToLogBook verifyAddedScanNote verifyNoteValue verifyLogBookDisplaysCurrentDate verifyLogbookDateBar verifyLogbookDateBarLeft clickCalendar verifyRequiredPreviousDateIsDisplayed pickAndVerifyPreviousDateIsDisplayed verifyLogbookDateBarRight clickCalanderRightButton VerifyRequiredDateIsDisplayed clickAddedNote verifyLogbookDetailDateTime | 1151
SDAUIRS881 | Click the Calendar icon in logbook & verify app displays the Calendar. Verify 89 days before the current day and ending on the current day is enabled in Calendar. Verify 90 days before the current day and days after current day are disabled in the calendar. Change the phone date & verify the above test case." Scan the sensor. Click the Calendar & select each day & verify app displays the list of scan results, time change events and notes for each day( if there is data collected for that day) in the period beginning 89 days before the current day and ending on the current day in Log book screen. Verify the information is displayed correctly in Logbook screen. Click on the Note & scan results and verify app displays the glucose graph in Logbook detail screen. Verify the information is displayed correctly. If there is no real time data on that day verify app displays 'No entries exist for this date' . | UILogbook_T007_Logbook_Glucose_Graph_for_90_days_Auto | navigateToLogBook clickCalendar verifyLeapYear verifyLogBookDisplaysStaticDate verifyAddedScanNote verifyNoteValue verifyCalenderDateDifference verifyLogBookDisplaysStaticDate verifyDataWithBoldLetters verifyRequiredPreviousDateIsDisplayed clickAndVerifyGlucoseGraph selectDateWithoutBoldLetter verifyLogBookEmpty | 1151
SDAUIRS885 | Verify added real time value is displayed correctly in graph. Verify in Logbook Detail Page. Verify the color code of real time point. Verify with different ranges of real time values(with in target, outside target range, LO & HI values) | UILogbook_T008_Logbook_Graph_Realtime_Historic_BG_scans | bgValueWithStaticTime scanMockSensor navigateToLogBook verifyandClickLogbookScandata Manual Image Verification | 1205 1221 1222 1223 1224 1225 1226
SDAUIRS886 | Add historic values. Scan Sensor. Verify added historic value is displayed correctly in graph in Logbook Detail Page. | UILogbook_T008_Logbook_Graph_Realtime_Historic_BG_scans | addMockSensorData scanMockSensor navigateToLogBook verifyandClickLogbookScandata Manual Image Verification | 1205 1228
SDAUIRS879* | Add a note & verify note information are displayed in Log book. Click on note in Log book & verify note information are displayed correctly in Log book Detail Page. Edit a existing note in Log book & verify note information are modified correctly in Log book Detail Page. | UILogbook_T009_Logbook_Notes_info_NonActionable_indication_and_HistoricBGScans_graph_below_70_mgdL | bgValueWithStaticTime scanMockSensor clickAddNote clickScanAddNoteButton selectAndVerifyFoodItemFromMenu setAndGetCarbAmount setAndVerifyInsulin selectExerciseAttribute clickExerciseIntensitySpinner setAndGetIntensity verifyExcerciseIntensityTimeDurationRanges clickDoneButton navigateToLogBook verifyAddedScanNote verifyScantime clickOnScanNote verifyMultipleNoteLogBookDetailPage verifyMultipleNoteDetailPage verifyScanAttributes verifyaddedNote verifyAddedNoteWithBGValue | 1176
SDAUIRS891 | Add a multiple notes at different time interval. Verify app displays the Note icon approximately in graph. Also click the note icon & verify app displays the note information & note time correctly. | UILogbook_T009_Logbook_Notes_info_NonActionable_indication_and_HistoricBGScans_graph_below_70_mgdL | clickAddedNote verifyTimeStamp navigateToLogBook selectAndVerifyFoodItemFromMenu setAndGetCarbAmount setAndVerifyInsulin selectExerciseAttribute clickExerciseIntensitySpinner setAndGetIntensity verifyExcerciseIntensityTimeDurationRanges clickDoneButton | 1176
SDAUIRS888 | Make sure two historic values within 20 mins are less than 70 mg/dl. Verify graph displays the connecting line in red. Verify historic values 69,71,70 mg/dl Verify > 70mg/dl shouldn't be in red color | UILogbook_T009_Logbook_Notes_info_NonActionable_indication_and_HistoricBGScans_graph_below_70_mgdL | bgValueWithStaticTime scanMockSensor navigateToLogBook clickOnScanNote Manual Image Verification | 1176 1229 1230
SDAUIRS887 | Add two or three historic values within 20 mins gap. Verify historic glucose values are connected. Add two or three historic values > 20 mins gap. Verify historic glucose values are not connected. Add two or three historic values exactly at 20 mins gap.Verify historic glucose values are connected. | UILogbook_T010_LogbookGraph_HistoricBGScans_In20Mins | bgValueWithStaticTime addMockSensorData scanMockSensor navigateToLogBook clickOnScanNote Manual Image Verification | 1177 1232 1233 1234
SDAUIRS890 | Change the timezone. Add historic data & scan the sensor. Change the timezone. Add historic data & scan the sensor. Verify graph doesn't show the historical glucose traces from before a time change event to overlap with those after a time change event. | UILogbook_T011_LogbookGraph_NotAllowHGonTimeChangeOverlap | navigateToLogBook clickCalendar pickAndVerifyPreviousDateIsDisplayed clickOnScanNote verifyTimeChangeIcon Manual Image Verification | 1178 1235 1236
SDAUIRS1188 | Verify app displays the unit of measure as mg/dL. Verify app allows to enter the values within the range ( 20 to 500) & app doesn't show any warning message. Verify app allows to enter the gram values boundary range ( 20 & 500) & app doesn't show any warning message. Verify app doesn't allow to enter the gram values outside the range ( > 500 & < 20) & app shows warning message. Verify app doesn't allow to enter decimal values. | ManualBG_T001_AddManualBGResult | setBGR clickBGRDone verifyBGError clearBGREntry | 1239
SDAUIRS909 | Verify app displays the unit of measure as mg/dL. Verify app allows to enter the values within the range ( 20 to 500) & app doesn't show any warning message. Verify app allows to enter the gram values boundary range ( 20 & 500) & app doesn't show any warning message. Verify app doesn't allow to enter the gram values outside the range ( > 500 & < 20) & app shows warning message. Verify app doesn't allow to enter decimal values. | ManualBG_T001_AddManualBGResult | clickBloodDropIcon setAndVerifyLogBookTime clickOnCursorBlink verifyOpenKeyBoard clearBGREntry setBGR clickBGRDone verifyLogBookBGR verifyNoBGError clickBGRCancel | 1239
SDAUIRS908 | Verify Manual BG icon is displayed in Logbook Page. Click on Manual BG icon & enter a BG value. Verify app allows to enter Manual BG successfully. | ManualBG_T001_AddManualBGResult | navigateToLogBook pickAndVerifyPreviousDateIsDisplayed verifyBloodDropIcon clickBloodDropIcon verifyBloodGlucoseResultPage clickBGRCancel verifyLogBookEmpty setAndVerifyLogBookTime setBGR clickBGRDone verifyaddedBGR | 1239
SDAUIRS1219 | Verify added Manual BG is displayed in Logbook screen. Verify BG Value in mg/dL, Manual BG icon , added time & timezone are displayed. | ManualBG_T001_AddManualBGResult | verifyLogBookBGR | 1239
SDAUIRS1220 | Scan a sensor & add a note Add a Manual BG. Verify app distinguish manually entered blood glucose values and notes. | ManualBG_T001_AddManualBGResult | addRealtimeMockSensor scanMockSensor navigateToLogBook clickBloodDropIcon setAndVerifyLogBookTime setBGR clickBGRDone verifyLogBookBGR verifyAddedScanNote verifyBlooddropNotFound | 1239
SDAUIRS1182 | Click on Manual BG icon. Verify app displays the timer to set time. Verify app allows to select the hours & minutes successfully . Note: Verify both 12 / 24 hr format timer Click the Calendar icon in Logbook screen & set the date. Click on Manual BG icon & set the time. Enter the valid BG value. Verify app creates the Manual BG is added successfully. Also verify the date, time,time zone, BG value matches." | ManualBG_T001_AddManualBGResult | verifyCalander navigateToLogBook clickBloodDropIcon verifyTimePickerHourTimeFormat | 1239
SDAUIRS259* | Add a Manual BG. Verify added Manual BG is displayed in Logbook screen by doing the below items. Restart the phone. Verify timepicker dialog doesn't allow to set in seconds. | ManualBG_T001_AddManualBGResult | verifyselectedLogBookTime clickTimepickerDone | 1239

## Methods - NOT Req Specific
Test Script Name | Methods - NOT Req Specific
ManualBG_T001_AddManualBGResult | changePhoneHourTimeFormat launch verifyLibreLinkLaunch deleteNoteIfAvailable clickCalendar pickDate verifyPickedDateIsDisplayed setLogBookTime setToDefault selectingSASMode clickBackArrowORMenu
UILogbook_T001_Logbook_Navigation | currentSystemTime changePhoneHourTimeFormat debugDrawerClearData clickAddNote setLogBookTime addNoteFoodOptions setAndVerifyInsulin selectExerciseAttribute clickDoneButton debugDrawerClearData setToDefault navigateToReminderScreen navigateToReports
UILogbook_T002_Logbook_Graph_Axis | setTheDateAndTime loadTestData changePhoneHourTimeFormat currentSystemTime setToDefault debugDrawerClearData
UILogbook_T004_Logbook_Scan_Note_Verification | setTargetGlucoseRange setTheTime changePhoneHourTimeFormat debugDrawerClearData openDebugDrawer scanMockSensor clickBackArrowORMenu clickCalendar pickAndVerifyPreviousDateIsDisplayed selectLogbookNote clickEditNote verifyNotesCommentsAvailability enterNotesComments clickDoneButton setToDefault debugDrawerClearData currentSystemTime
UILogbook_T005_Logbook_Change_Time_Zone | changePhoneHourTimeFormat setTheTime launchAppWithdata debugDrawerClearData selectingSASMode closeDebugDrawer bgValueWithStaticTime clickBackArrowORMenu timezoneCleanUp appNetworkTimeCheck changeTimeZone setAdvanceStaticDateAndTime currentSystemTime
UILogbook_T006_Logbook_Graph_Time_Zone_changes | appNetworkTimeCheck setTheTime selectingSASMode launchAppWithdata debugDrawerClearData clickBackArrowORMenu closeDebugDrawer verifyLogbookScanDetail timezoneCleanUp changeTimeZone currentSystemTime changePhoneHourTimeFormat
UILogbook_T007_Logbook_Glucose_Graph_for_90_days_Auto | changePhoneHourTimeFormat setTheTime setTheDateAndTime loadTestData appNetworkTimeCheck clickBackArrowORMenu setToDefault currentSystemTime debugDrawerClearData getDeviceDate
UILogbook_T008_Logbook_Graph_Realtime_Historic_BG_scans | changePhoneHourTimeFormat setTheTime debugDrawerClearData setToDefault selectingSASMode setTargetGlucoseRange openDebugDrawer clickBackArrowORMenu advanceDeviceTime closeDebugDrawer setToDefault currentSystemTime bgValueWithStaticTime
UILogbook_T009_Logbook_Notes_info_NonActionable_indication_and_HistoricBGScans_graph_below_70_mgdL | appNetworkTimeCheck changePhoneHourTimeFormat debugDrawerClearData openDebugDrawer selectingSASMode setAndGetCarbUnit setTargetGlucoseRange clickBackArrowORMenu doSwipeDown closeBalloon closeDebugDrawer navigateToHomePage advanceDeviceTime changeTimeZone setTheTime launchAppWithdata timezoneCleanUp currentSystemTime
UILogbook_T010_LogbookGraph_HistoricBGScans_In20Mins | changePhoneHourTimeFormat setTheTime selectingSASMode clickBackArrowORMenu setToDefault currentSystemTime debugDrawerClearData openDebugDrawer
UILogbook_T011_LogbookGraph_NotAllowHGonTimeChangeOverlap | loadTestData timezoneCleanUp currentSystemTime changePhoneHourTimeFormat appNetworkTimeCheck changeTimeZone launchAppWithdata setTheTime getDeviceDate