## Traceability Matrix
Req ID | Test Cases | Test Script Name | Script Method Name | DVT Protocol Test Step ID
SDASRS1164 | Verify Time In Target report is displayed under Reports. Verify Daily Graph report is displayed under Reports. | DailyPatterns_T001_ViewReport_DisplayAGP_ReportPeriods_Xaxis_targetRange | navigateToReports navigateToReportBySlide verifyDailyPatternsReport | 1237
SDASRS1169 | Verify Daily Graph report is displayed under Reports. | DailyGraph_T001_View_Display_Xaxis_Distinguish_data_Notes | navigateToReports navigateToReportBySlide verifyDailyGraphReport | 1150
SDAUIRS1003 | Verify Daily Graph report is displayed under Reports. | DailyGraph_T001_View_Display_Xaxis_Distinguish_data_Notes | navigateToReports navigateToReportBySlide verifyDailyGraphReport | 1150
SDAUIRS1004 | Verify app allows us to select & view each day report in the last 90 days. Verify app displays the report for each day correctly as per the dataset. Verify app allows us to select & view each day report in the last 90 days. Verify other days other than 90 days are disabled in the calendar. | DailyGraph_T001_View_Display_Xaxis_Distinguish_data_Notes | pickDate clickOnCalendarDate clickArrowOnTopRightOfTheGraph clickArrowOnTopLeftOfTheGraph verifyDailyGraphViewDate90Days verifyDailyGraphViewDate89Days verifyDailyGraphViewDate91Day verifyDailyGraphViewDateFutureDay verifyDailyGraphViewDateCurrentDay verifyDailyGraphLeftArrowWith10Times verifyDailyGraphViewDateDay verifyDailyGraphReport verifyDailyGraphRightArrowWith5Times | 1150
SDAUIRS1005 | Verify app displays the glucose graph. Verify app displays the graph for each day correctly as per the dataset. | DailyGraph_T001_View_Display_Xaxis_Distinguish_data_Notes | verifyDailyGraphReport | 1150
SDAUIRS1195 | Verify graph distinguishes the target glucose range. | DailyGraph_T001_View_Display_Xaxis_Distinguish_data_Notes | navigateToReports verifyDailyGraphReport Manual Image Verification | 1150 1303 1304
SDAUIRS1006 | Verify graph displays the X axis range of midnight to midnight. Change the Phone's time format as 24 hr. Verify graph displays the X axis range of midnight to midnight. | DailyGraph_T001_View_Display_Xaxis_Distinguish_data_Notes | verifyReportXAxisWith24HoursFormat verifyReportXAxisWith12HoursFormat | 1150
SDAUIRS1011 | Verify graph displays the connecting line in red. Verify historic values 69,71,70 mg/dl Verify > =70mg/dl shouldn't be in red color | DailyGraph_T001_View_Display_Xaxis_Distinguish_data_Notes | verifyDailyGraphReport Manual Image Verification pickDate | 1150 1305 1306
SDAUIRS1014 | Verify graph displays the badge indicating the number of notes being represented by that icon. Verify the number of notes is displayed correctly. Verify the note icon time plotted in graph matches approximately with the note time. | DailyGraph_T001_View_Display_Xaxis_Distinguish_data_Notes | navigateToReports navigateToLogBook pickDate clickOnCalendarDate add3Notesat5amToLogbook add5NotesatNoonToLogbook add4Notesat2amToLogbook selectingNotes verifyNote Manual Image Verification | 1150 1307 1308 1309 1311 1312 1313
SDAUIRS1007 | Verify graph displays the Y-axis from 0 to at least 350 mg/dL. (0-21 mmol/L) Verify graph displays the Y-axis from 0 to 500 mg/dL if glucose data > 350 mg/dL (0- 27 mmol/L) | DailyGraph_T002_Gap_LowG_YaxisRange | navigateToReports clickArrowOnTopLeftOfTheGraph | 1205 1314 1315 1316
SDAUIRS1009 | Verify app displays the historic glucose values in the graph correctly as per the dataset. | DailyGraph_T002_Gap_LowG_YaxisRange | Manual Image Verification | 1205 1317
SDAUIRS1010 | Verify historic glucose values are connected if gap < 20mins. Verify historic glucose values are not connected if gap is > 20mins. Verify historic glucose values are connected if gap is exactly at 20 mins | DailyGraph_T002_Gap_LowG_YaxisRange | Manual Image Verification | 1205 1318 1319 1320
SDAUIRS1012 | Add a historic data. Scan the sensor Change the current timezone to different timezone. Add a historic data. Scan the sensor. Verify timechange icon is displayed. Also Verify the timechange icon is plotted correctly. | DailyGraph_T003_TimeChange | navigateToReports clickArrowOnTopLeftOfTheGraph verifyTimeZoneIcon Manual Image Verification | 1176 1322
SDAUIRS1013 | Verify graph doesn't show the historical glucose traces from before a time change event to overlap with those after a time change event. | DailyGraph_T003_TimeChange | Manual Image Verification | 1176 1321 1323
SDAUIRS961 | Verify Daily patterns report is displayed under Reports. | DailyPatterns_T001_ViewReport_DisplayAGP_ReportPeriods_Xaxis_targetRange | navigateToReports navigateToReportBySlide verifyDailyPatternsReport | 1237
SDAUIRS962 | Load the dataset. Select 7 days/14 days /30 days/90 days. Verify displayed Daily patterns graph,data available & percentiles are matches with the Apollo Daily patterns Report. | DailyPatterns_T001_ViewReport_DisplayAGP_ReportPeriods_Xaxis_targetRange | navigateToReports clickOnDays Manual Image Verification | 1237 1324 1329 1330 1331
SDAUIRS963 | Verify app provides the reporting periods as 7 days, 14 days, 30 days, 90 days. Select each reporting period & verify app displays the report for each period correctly as per the dataset. | DailyPatterns_T001_ViewReport_DisplayAGP_ReportPeriods_Xaxis_targetRange | navigateToReports clickOnDays verifyDailyPatternDatePeriodForNDays | 1237
SDAUIRS964 | Change the Phone's time format as 12 hr. Verify graph displays the X axis range of midnight to midnight. Change the Phone's time format as 24 hr. Verify graph displays the X axis range of midnight to midnight. | DailyPatterns_T001_ViewReport_DisplayAGP_ReportPeriods_Xaxis_targetRange | navigateToReports clickOnDays verifyReportXAxisWith24HoursFormat verifyReportXAxisWith12HoursFormat | 1237
SDAUIRS966 | Verify graph distinguishes the target glucose range. Change the target glucose range and verify that the new range is displayed correctly on the report. | DailyPatterns_T001_ViewReport_DisplayAGP_ReportPeriods_Xaxis_targetRange | navigateToReports clickOnDays Manual Image Varification | 1237 1332 1333 1334 1335 1336 1337 1338 1339
SDASRS1164 | Verify Time In Target report is displayed under Reports. Click on Time In Target report & verify Time In Target report is displayed | DailyPatterns_T001_ViewReport_DisplayAGP_ReportPeriods_Xaxis_targetRange | navigateToReports navigateToReportBySlide verifyDailyPatternsReport | 1237
SDASRS1165 | Verify Daily patterns report is displayed under Reports. Click on Daily patterns report & verify Daily patterns report is displayed | DailyPatterns_T001_ViewReport_DisplayAGP_ReportPeriods_Xaxis_targetRange | navigateToReports navigateToReportBySlide verifyDailyPatternsReport clickOnDays | 1237
SDAUIRS965 | Verify graph displays the Y axis range of 0 to at least 350 mg/dL.(0 -21 mmol/L) Verify graph displays the Y axis range of 0 to 500 mg/dL if glucose data is > 350 (0 -27 mmol/L) | DailyPatterns_T002_Yaxis | navigateToReports clickOnDays Manual Image Verification | 1238 1340 1343 1341 1342 1344 1345 1346 1347
SDAUIRS967 | Verify app displays the number of days within the reporting period that contain at least one historical glucose point. Load the dataset with valid historic glucose points & verify app displays the number of days correctly as per the dataset. Verify app displays the number of days correctly and matches with the dataset. | DailyPatterns_T003_DataAvailable | navigateToReports clickOnDays verifyDailyPatternReportDataAvailableForNdays | 1239
SDAUIRS956 | Load the dataset. Select 7 days/14 days /30 days/90 days. Verify displayed Daily patterns graph & percentiles are matches with the Apollo Daily patterns Report. | DailyPatterns_T004_S24Req_Test | navigateToReports clickOnDays Manual Image Verification | 1241 1348 1349 1350
SDAUIRS941 | Load the dataset. Select 7 days/14 days /30 days/90 days. Verify displayed Daily patterns graph & percentiles are matches with the Apollo Daily patterns Report. | DailyPatterns_T004_S24Req_Test | navigateToReports clickOnDays Manual Image Verification | 1241 1348 1349 1350
SDAUIRS928 | Load the dataset. Select 7 days/14 days /30 days/90 days. Verify displayed Daily patterns matches with the Apollo Daily patterns Report. | DailyPatterns_T004_S24Req_Test | navigateToReports clickOnDays Manual Image Verification | 1241 1348 1349 1350
SDAUIRS1016 | Verify Estimated A1c report is displayed under Reports. Click on Estimated A1c report & verify Estimated A1c report is displayed | A1c_T001_View_Display_calculation | navigateToReports navigateToReportBySlide verifyEstimatedA1cReport | 1243
SDAUIRS1017 | Verify that we are using 89 actual days and not using months. | A1c_T001_View_Display_calculation | navigateToReports verifyEstimatedA1cReportDatePeriodForNdays | 1243
SDAUIRS1018 | Verify app displays the number of days including and between the first point and the last point within the reporting period. Load the dataset & verify app displays the number of days correctly as per the dataset. | A1c_T001_View_Display_calculation | verifyEstimatedA1cReportDataSpanForNdays | 1243
SDAUIRS1249 | Verify the app displays the Estimated A1c correctly as per the dataset & as per the algorithm defined in DOC29452. | A1c_T001_View_Display_calculation | verifyEstimatedA1cReportCalPercentageForNdays verifyEstimatedA1cReportCalmmolForNdays | 1243
SDASRS1186 | Verify Estimated A1c report is displayed under Reports. Click on Estimated A1c report & verify Estimated A1c report is displayed | A1c_T001_View_Display_calculation | navigateToReports navigateToReportBySlide verifyEstimatedA1cReport | 1243
SDASRS1164 | Verify Low Glucose Events report is displayed under Reports. Click on Low Glucose Events report & verify Low Glucose Events report is displayed" Verify Average Glucose report is displayed under Reports. Click on Average Glucose report & verify Average Glucose report is displayed | DailyPatterns_T001_ViewReport_DisplayAGP_ReportPeriods_Xaxis_targetRange LowGlucoseEvents_T001_View_Display_Date_Xaxis A1c_T001_View_Display_calculation DailyGraph_T001_View_Display_Xaxis_Distinguish_data_Notes | navigateToReports navigateToReportBySlide verifyDailyPatternsReport verifyEstimatedA1cReport verifyDailyGraphReport verifyLowGlucoseEventsReport | 1150 1243 1237 1244
SDAUIRS1017 | Verify that we are using 89 actual days and not using months | A1c_T002_View_Display | navigateToReports verifyEstimatedA1cDateHeader | 1242
SDAUIRS1018 | Verify app displays the number of days including and between the first point and the last point within the reporting period Load the dataset & verify app displays the number of days correctly as per the dataset. | A1c_T002_View_Display | navigateToReports verifyEstimatedA1cReportDataSpanForNdays | 1242
SDAUIRS1250 | Verify app displays the Estimated A1c as a percentage and in mmol/mol if the AGP calculated for the reporting period has at least 20 hours of data. Verify boundary conditions here (<20, 20, >20) | A1c_T002_View_Display | navigateToReports verifyNotEnounghSensorData verifyEstimatedA1cReportCalculatedValues | 1242 1363 1364
SDAUIRS979 | Verify Low Glucose Events report is displayed under Reports. Click on Low Glucose Events report & verify Low Glucose Events report is displayed | LowGlucoseEvents_T001_View_Display_Date_Xaxis | navigateToReports navigateToReportBySlide verifyLowGlucoseEventsReport | 1244
SDAUIRS981 | Verify app provides the reporting periods as 7 days, 14 days, 30 days, 90 days. Select each reporting period & verify app displays the report for each period correctly as per the dataset. | LowGlucoseEvents_T001_View_Display_Date_Xaxis | navigateToReports clickOnDays verifyLowGlucoseEventsDatePeriodForNDays | 1244
SDAUIRS985 | Change the Phone's time format as 12 hr. Verify graph displays the X axis range of midnight to midnight. Change the Phone's time format as 24 hr. Verify graph displays the X axis range of midnight to midnight. | LowGlucoseEvents_T001_View_Display_Date_Xaxis | navigateToReports clickOnDays verifyReportXAxisWith24HoursFormat verifyReportXAxisWith12HoursFormat | 1244
SDAUIRS987 | Verify app displays one bar per 3-hour period, beginning at midnight in the graph. | LowGlucoseEvents_T001_View_Display_Date_Xaxis | navigateToReports clickOnDays verifyReportXAxisWith3HoursPeriodFor24Hours verifyReportXAxisWith3HoursPeriodFor12Hours | 1244
SDASRS1164 | Click on Daily patterns report & verify Daily patterns report is displayed Click on Time In Target report & verify Time In Target report is displayed | DailyPatterns_T001_ViewReport_DisplayAGP_ReportPeriods_Xaxis_targetRange | navigateToReports navigateToReportBySlide verifyDailyPatternsReport | 1237
SDASRS1167 | Verify Low Glucose Events report is displayed under Reports. Click on Low Glucose Events report & verify Low Glucose Events report is displayed | LowGlucoseEvents_T001_View_Display_Date_Xaxis | navigateToReports navigateToReportBySlide verifyLowGlucoseEventsReport | 1244
SDAUIRS983 | Verify app displays the number of days within the reporting period that contain at least one historical glucose point. Load the dataset & verify app displays the number of days correctly as per the dataset. Verify app displays the number of days correctly and matches with the dataset. | LowGlucoseEvents_T002_DataAvailable | navigateToReports clickOnDays verifyLowGlucoseEventsReportDataAvailabilityForNdays | 1246
SDAUIRS982 | Load the dataset & verify app displays the total number of low glucose events that began during the reporting period. Verify for all reporting periods. | LowGlucoseEvents_T003_Calculation_Yaxis | navigateToReports clickOnDays verifyLowGlucoseEventsReportForTotalEvents | 1245
SDAUIRS984 | Load the dataset & verify app displays the bar graph showing the number of low glucose events that began in each time period of the day. Verify for all reporting periods. | LowGlucoseEvents_T003_Calculation_Yaxis | navigateToReports clickOnDays Image Verification Manually | 1245 1352 1353 1354 1355 1356 1357 1358
SDAUIRS986 | Verify app displays the Y-axis whose range is from 0 to at least the highest bar contained within the graph. Verify the Y axis is not labelled. Verify the Y axis for different ranges. | LowGlucoseEvents_T003_Calculation_Yaxis | navigateToReports clickOnDays Image Verification Manually | 1245 1359 1360 1361 1362
SDAUIRS989 | Verify app displays the count of events corresponding to each bar. Load the dataset & verify app displays the count correctly as per the dataset. | LowGlucoseEvents_T003_Calculation_Yaxis | navigateToReports clickOnDays Image Verification Manually | 1245 1352 1353 1354 1355 1356 1357 1358

## Methods - NOT Req Specific
Test Script Name | Methods - NOT Req Specific
A1c_T001_View_Display_calculation | changePhoneHourTimeFormat setTheDateAndTime loadTestData setTheDate setToDefault navigateToHomePage navigateToLogBook navigateToReminder currentSystemTime debugDrawerClearData
A1c_T002_View_Display | changePhoneHourTimeFormat setTheDateAndTime loadTestData setTheDate setToDefault currentSystemTime debugDrawerClearData navigateToHomePage
DailyGraph_T001_View_Display_Xaxis_Distinguish_data_Notes | changePhoneHourTimeFormat setTheDateAndTime loadTestData setTargetGlucoseRange navigateToHomePage clickOnDone clickOnGlucoseUnits moveToLogbook clickOnLogbookDetailAddNote selectingNotes setLogBookTime clickAddNote setToDefault navigateToReminder navigateToLogBook currentSystemTime debugDrawerClearData
DailyGraph_T002_Gap_LowG_YaxisRange | changePhoneHourTimeFormat setTheDateAndTime loadTestData navigateToHomePage setToDefault setTargetGlucoseRange currentSystemTime debugDrawerClearData
DailyGraph_T003_TimeChange | setToDefault changeTimeZone changePhoneHourTimeFormat setTheDateAndTime launchAppWithdata loadTestData automaticDateAndTime automaticDateAndTimeZone appNetworkTimeCheck currentSystemTime debugDrawerClearData
DailyPatterns_T001_ViewReport_DisplayAGP_ReportPeriods_Xaxis_targetRange | changePhoneHourTimeFormat setTheDateAndTime loadTestData setTargetGlucoseRange setToDefault navigateToHomePage navigateToLogBook navigateToReminder currentSystemTime debugDrawerClearData
DailyPatterns_T002_Yaxis | changePhoneHourTimeFormat setTheDateAndTime loadTestData navigateToHomePage setToDefault currentSystemTime debugDrawerClearData
DailyPatterns_T003_DataAvailable | changePhoneHourTimeFormat setTheDateAndTime loadTestData setToDefault currentSystemTime debugDrawerClearData
DailyPatterns_T004_S24Req_Test | changePhoneHourTimeFormat setTheDate loadTestData appNetworkTimeCheck navigateToHomePage setTargetGlucoseRange setToDefault currentSystemTime debugDrawerClearData
LowGlucoseEvents_T001_View_Display_Date_Xaxis | changePhoneHourTimeFormat setTheDateAndTime loadTestData setToDefault navigateToHomePage navigateToLogBook navigateToReminder currentSystemTime debugDrawerClearData
LowGlucoseEvents_T002_DataAvailable | changePhoneHourTimeFormat setTheDateAndTime loadTestData setToDefault currentSystemTime debugDrawerClearData
LowGlucoseEvents_T003_Calculation_Yaxis | changePhoneHourTimeFormat setTheDateAndTime loadTestData navigateToHomePage setToDefault currentSystemTime debugDrawerClearData