## Traceability Matrix
Req ID | Test Cases | Test Script Name | Script Method Name | DVT Protocol Test Step ID
SDASRS1164 | Verify Sensor Usage report is displayed under Reports. Verify Sensor Usage report is displayed. Verify Low Glucose Events report is displayed under Reports. verify Low Glucose Events report is displayed. | AverageGlucose_T001_View_Display_Xaxis | navigateToReports verifyAverageGlucoseReport | 1208
SDASRS1168 | Verify Average Glucose report is displayed under Reports. Verify Average Glucose report is displayed | AverageGlucose_T001_View_Display_Xaxis | navigateToReports verifyAverageGlucoseReport | 1208
SDAUIRS1000 | Verify app displays average glucose corresponding to each bar correctly as per the dataset. | AverageGlucose_T003_Calculation_Distinguish_Data_mg AverageGlucose_T004_Calculation_Distinguish_Data_mmol | navigateToReports clickOnDays Manual Verification | 1248 1250 1251 1252 1253 1260 1262 1266 1269 1149 1274 1275 1276 1277 1285 1292 1296 1297
SDAUIRS1196 | Verify app distinguish the average historical glucose readings that are < 70, with in target range, greater than or equal to 70 mg/dL and less than the lower target., ess than or equal to 240 mg/dL and greater than the upper target,greater than 240 mg/dL. in the report as per the dataset. Verify for all reporting periods with different target ranges. Verify the color of bars are displayed correctly. | AverageGlucose_T003_Calculation_Distinguish_Data_mg AverageGlucose_T004_Calculation_Distinguish_Data_mmol | navigateToReports clickOnDays Manual Verification | 1248 1389 1395 1398 1401 1149 1416 1419 1422
SDAUIRS1197 | Verify app distinguish the average historical glucose readings that are < 70, with in target range, greater than or equal to 70 mg/dL and less than the lower target., ess than or equal to 240 mg/dL and greater than the upper target,greater than 240 mg/dL. in the report as per the dataset. Verify for all reporting periods with different target ranges. Verify the color of bars are displayed correctly. | AverageGlucose_T003_Calculation_Distinguish_Data_mg AverageGlucose_T004_Calculation_Distinguish_Data_mmol | navigateToReports clickOnDays Manual Verification | 1248 1369 1373 1377 1381 1390 1396 1399 1402 1149 1405 1408 1411 1414 1417 1420 1423 1426
SDAUIRS1198 | Verify app distinguish the average historical glucose readings that are < 70, with in target range, greater than or equal to 70 mg/dL and less than the lower target., ess than or equal to 240 mg/dL and greater than the upper target,greater than 240 mg/dL. in the report as per the dataset. Verify for all reporting periods with different target ranges. Verify the color of bars are displayed correctly. | AverageGlucose_T003_Calculation_Distinguish_Data_mg AverageGlucose_T004_Calculation_Distinguish_Data_mmol | navigateToReports clickOnDays Manual Verification | 1248 1261 1386 1387 1391 1397 1400 1403 1149 1404 1407 1410 1413 1418 1421 1424 1427
SDAUIRS1199 | Verify app distinguish the average historical glucose readings that are < 70, with in target range, greater than or equal to 70 mg/dL and less than the lower target., ess than or equal to 240 mg/dL and greater than the upper target,greater than 240 mg/dL. in the report as per the dataset. Verify for all reporting periods with different target ranges. Verify the color of bars are displayed correctly. | AverageGlucose_T003_Calculation_Distinguish_Data_mg AverageGlucose_T004_Calculation_Distinguish_Data_mmol | navigateToReports clickOnDays Manual Verification | 1248 1370 1385 1375 1379 1383 1149 1406 1409 1412 1415
SDAUIRS1200 | Verify app distinguish the average historical glucose readings that are < 70, with in target range, greater than or equal to 70 mg/dL and less than the lower target., ess than or equal to 240 mg/dL and greater than the upper target,greater than 240 mg/dL. in the report as per the dataset. Verify for all reporting periods with different target ranges. Verify the color of bars are displayed correctly. | AverageGlucose_T003_Calculation_Distinguish_Data_mg AverageGlucose_T004_Calculation_Distinguish_Data_mmol | navigateToReports clickOnDays Manual Verification | 1248 1261 1385 1386 1387 1149 1404 1407 1410 1413
SDAUIRS991 | Verify Average Glucose report is displayed under Reports. Verify Average Glucose report is displayed | AverageGlucose_T001_View_Display_Xaxis | navigateToReports navigateToReportBySlide verifyAverageGlucoseReport | 1208
SDAUIRS992 | Verify app provides the reporting periods as 7 days, 14 days, 30 days, 90 days. Verify app displays the report for each period correctly as per the dataset. | AverageGlucose_T001_View_Display_Xaxis | navigateToReports clickOnDays verifyAverageGlucoseDatePeriodForNDays | 1208
SDAUIRS993 | Verify app displays the average value of all historical glucose readings in the reporting period. Verify the displayed average value matches with the dataset. Verify the UOM is displayed. | AverageGlucose_T003_Calculation_Distinguish_Data_mg AverageGlucose_T004_Calculation_Distinguish_Data_mmol | navigateToReports verifyAverageHistoricGlucoseForNdays verifyAverageHistoricGlucoseRange155To180For7Days verifyTotalAverageHistoricalGlucoseReadingFor7Days verifyLowAverageHistoricalGlucoseReadingValuesFor7Days clickOnDays | 1248 1149 1280
SDAUIRS994 | Verify app displays the number of days within the reporting period that contain at least one historical glucose point. Verify app displays the number of days correctly as per the dataset. Verify app displays the number of days correctly and matches with the dataset. | AverageGlucose_T002_DataAvailable | navigateToReports clickOnDays verifyAverageGlucoseDataAvailableForNdays | 1147
SDAUIRS995 | Verify app displays bar graph showing the average of all historical glucose readings within each time period of the day correctly as per the dataset. Verify bars are distinguished correctly. | AverageGlucose_T003_Calculation_Distinguish_Data_mg AverageGlucose_T004_Calculation_Distinguish_Data_mmol | navigateToReports clickOnDays Manual Verification | 1248 1250 1251 1252 1253 1254 1149 1278
SDAUIRS996 | Verify graph displays the X axis range of midnight to midnight. Verify graph displays the X axis range of midnight to midnight. | AverageGlucose_T001_View_Display_Xaxis | clickOnDays verifyReportXAxisWith24HoursFormat verifyReportXAxisWith12HoursFormat | 1208
SDAUIRS997 | Verify app displays the Y-axis whose range is from 0 to at least the highest bar contained within the graph. Verify the Y axis is not labelled. Verify the Y axis for different ranges. | AverageGlucose_T003_Calculation_Distinguish_Data_mg AverageGlucose_T004_Calculation_Distinguish_Data_mmol | navigateToReports clickOnDays Manual Verification | 1248 1255 1256 1257 1258 1264 1265 1149 1274 1275 1276 1277 1281 1282 1283 1284 1494 1495
SDAUIRS998 | Verify app displays one bar per 3-hour period, beginning at midnight in the graph. | AverageGlucose_T001_View_Display_Xaxis | clickOnDays verifyReportXAxisWith3HoursPeriodFor24Hours verifyReportXAxisWith3HoursPeriodFor12Hours | 1208

## Methods - NOT Req Specific
Test Script Name | Methods - NOT Req Specific
AverageGlucose_T001_View_Display_Xaxis | changePhoneHourTimeFormat setTheDateAndTime loadTestData navigateToHomePage navigateToLogBook navigateToReminder setToDefault debugDrawerClearData navigateToReportBySlide
AverageGlucose_T002_DataAvailable | changePhoneHourTimeFormat setTheDateAndTime loadTestData navigateToHomePage setToDefault currentSystemTime debugDrawerClearData
AverageGlucose_T003_Calculation_Distinguish_Data_mg AverageGlucose_T004_Calculation_Distinguish_Data_mmol | changePhoneHourTimeFormat setTheDateAndTime loadTestData navigateToHomePage setTargetGlucoseRange setToDefault debugDrawerClearData currentSystemTime