## Traceability Matrix
Req ID | Test Cases | Test Script Name | Script Method Name | DVT Protocol Test Step ID
SDASRS1164 | Verify Daily patterns report is displayed under Reports. Click on Daily patterns report & verify Daily patterns report is displayed Verify Time In Target report is displayed under Reports. Click on Time In Target report & verify Time In Target report is displayed | SensorUsage_T001_View_Display TimeInTarget_T001_ViewReport_ReportPeriods_TargetRange | navigateToReports navigateToReportBySlide verifySensorUsageReport verifyTimeInTargetReport | 1205 1177
SDASRS1166 | Verify Time In Target report is displayed under Reports. Click on Time In Target report & verify Time In Target report is displayed | TimeInTarget_T001_ViewReport_ReportPeriods_TargetRange | navigateToReports verifyTimeInTargetReport navigateToReportBySlide | 1177
SDASRS1170 | Verify Sensor Usage report is displayed under Reports. Click on Sensor Usage report & verify Sensor Usage report is displayed | SensorUsage_T001_View_Display | navigateToReports navigateToReportBySlide verifySensorUsageReport | 1205
SDAUIRS1020 | Verify Sensor Usage report is displayed under Reports. Click on Sensor Usage report & verify Sensor Usage report is displayed | SensorUsage_T001_View_Display | navigateToReports navigateToReportBySlide verifySensorUsageReport | 1205
SDAUIRS1021 | Verify app provides the reporting periods as 7 days, 14 days, 30 days, 90 days. Select each reporting period & verify app displays the report for each period correctly as per the dataset. | SensorUsage_T001_View_Display | navigateToReports clickOnDays verifySensorGlucoseDatePeriodForNDays verifySensorUsageDateFor14DaysWithAdditionalTwoDays verifySensorUsageDateFor14DaysWithLeapYear | 1205
SDAUIRS1022 | Verify app displays the total number of realtime glucose values in the reporting period correctly as per the dataset. Verify app displays the average number of realtime glucose values obtained per day over the reporting period correctly as per the dataset. Verify app displays the percent of historical glucose data captured during the reporting period correctly as per the dataset. | SensorUsage_T002_Calculation | navigateToReports clickOnDays verifySensorUsageScanAttemptsForBasicData verifySensorUsageScanAttemptsMock1ErrorHITemp | 1176
SDAUIRS1023 | Verify app displays the total number of realtime glucose values in the reporting period correctly as per the dataset. Verify app displays the average number of realtime glucose values obtained per day over the reporting period correctly as per the dataset. Verify app displays the percent of historical glucose data captured during the reporting period correctly as per the dataset. | SensorUsage_T002_Calculation | clickOnDays verifySensorUsageAverageScansPerDayForBasicData verifySensorUsageAverageScansForMock1 | 1176
SDAUIRS1024 | Verify app displays the total number of realtime glucose values in the reporting period correctly as per the dataset. Verify app displays the average number of realtime glucose values obtained per day over the reporting period correctly as per the dataset. Verify app displays the percent of historical glucose data captured during the reporting period correctly as per the dataset. | SensorUsage_T002_Calculation | clickOnDays verifySensorUsagePercentageOfHGDataForBasicData verifySensorUsageHGPercentageForMock1 | 1176
SDAUIRS1193 | Verify graph distinguishes the current target glucose range. Change the target range and verify that the new range is displayed correctly on the report. Note: Verify for all reporting periods with different target ranges | TimeInTarget_T001_ViewReport_ReportPeriods_TargetRange | navigateToReports clickOnDays verifyTimeInTargetRange100To140 verifyTimeInTargetRange70To180 Manual Image Verification | 1177
SDAUIRS1194 | Load the dataset. Verify app displays the percentage of historical glucose points in the reporting period that are < 70, with in target range, greater than or equal to 70 mg/dL and less than the lower target., less than or equal to 240 mg/dL and greater than the upper target,greater than 240 mg/dL. in the report as per the dataset. Verify for all reporting periods with different target ranges. Verify the color of bars are displayed correctly. | TimeInTarget_T003_Percentage_DistinguishTargetRange_mg | navigateToReports clickOnDays Manual Image Verification | 1252 1280 1281 1282 1283
SDAUIRS1194 | Load the dataset. Verify app displays the percentage of historical glucose points in the reporting period that are < 70, with in target range, greater than or equal to 70 mg/dL and less than the lower target., less than or equal to 240 mg/dL and greater than the upper target,greater than 240 mg/dL. in the report as per the dataset. Verify for all reporting periods with different target ranges. Verify the color of bars are displayed correctly. | TimeInTarget_T004_Percentage_DistinguishTargetRange_mmol | navigateToReports clickOnDays Manual Image Verification | 1253 1305 1306 1307 1308
SDAUIRS969 | Verify Time In Target report is displayed under Reports. Click on Time In Target report & verify Time In Target report is displayed | TimeInTarget_T001_ViewReport_ReportPeriods_TargetRange | navigateToReports verifyTimeInTargetReport navigateToReportBySlide | 1177
SDAUIRS970 | Verify app provides the reporting periods as 7 days, 14 days, 30 days, 90 days. Select each reporting period & verify app displays the report for each period correctly as per the dataset. | TimeInTarget_T001_ViewReport_ReportPeriods_TargetRange | navigateToReports clickOnDays verifyTimeInTargetDatePeriodForNDays | 1177
SDAUIRS971 | Verify app displays the number of days within the reporting period that contain at least one historical glucose point. Load the dataset with valid historic glucose points & verify app displays the number of days correctly as per the dataset. | TimeInTarget_T002_DataAvailable | navigateToReports clickOnDays verifyTimeInTargetReportDataAvailabilityForNdays | 1251
SDAUIRS972 | Verify app visually distinguishes each of the displayed percentages. Verify the color of bars are displayed correctly. | TimeInTarget_T003_Percentage_DistinguishTargetRange_mg | navigateToReports clickOnDays Manual Image Verification | 1252 1265 1266 1267 1269 1270 1272 1273 1276 1277 1278 1279 1280 1281 1282 1283
SDAUIRS972 | Verify app visually distinguishes each of the displayed percentages. Verify the color of bars are displayed correctly. | TimeInTarget_T004_Percentage_DistinguishTargetRange_mmol | navigateToReports clickOnDays Manual Image Verification | 1253 1290 1291 1292 1294 1295 1297 1298 1301 1302 1303 1304 1305 1306 1307 1308
SDAUIRS973 | Load the dataset. Verify app displays the percentage of historical glucose points in the reporting period that are < 70, with in target range, greater than or equal to 70 mg/dL and less than the lower target., ess than or equal to 240 mg/dL and greater than the upper target,greater than 240 mg/dL. in the report as per the dataset. Verify for all reporting periods with different target ranges. Verify the color of bars are displayed correctly. | TimeInTarget_T003_Percentage_DistinguishTargetRange_mg | navigateToReports clickOnDays Manual Image Verification | 1252 1268 1269 1270 1271
SDAUIRS973 | Load the dataset. Verify app displays the percentage of historical glucose points in the reporting period that are < 70, with in target range, greater than or equal to 70 mg/dL and less than the lower target., less than or equal to 240 mg/dL and greater than the upper target,greater than 240 mg/dL. in the report as per the dataset. Verify for all reporting periods with different target ranges. Verify the color of bars are displayed correctly. | TimeInTarget_T004_Percentage_DistinguishTargetRange_mmol | navigateToReports clickOnDays Manual Image Verification | 1253 1293 1294 1295 1296
SDAUIRS974 | Load the dataset. Verify app displays the percentage of historical glucose points in the reporting period that are < 70, with in target range, greater than or equal to 70 mg/dL and less than the lower target., less than or equal to 240 mg/dL and greater than the upper target,greater than 240 mg/dL. in the report as per the dataset. Verify for all reporting periods with different target ranges. Verify the color of bars are displayed correctly. | TimeInTarget_T003_Percentage_DistinguishTargetRange_mg | navigateToReports clickOnDays Manual Image Verification | 1252 1264 1265 1266 1267
SDAUIRS974 | Load the dataset. Verify app displays the percentage of historical glucose points in the reporting period that are < 70, with in target range, greater than or equal to 70 mg/dL and less than the lower target., less than or equal to 240 mg/dL and greater than the upper target,greater than 240 mg/dL. in the report as per the dataset. Verify for all reporting periods with different target ranges. Verify the color of bars are displayed correctly. | TimeInTarget_T004_Percentage_DistinguishTargetRange_mmol | navigateToReports clickOnDays Manual Image Verification | 1253 1289 1290 1291 1292
SDAUIRS975 | Load the dataset. Verify app displays the percentage of historical glucose points in the reporting period that are < 70, with in target range, greater than or equal to 70 mg/dL and less than the lower target., less than or equal to 240 mg/dL and greater than the upper target,greater than 240 mg/dL. in the report as per the dataset. Verify for all reporting periods with different target ranges. Verify the color of bars are displayed correctly. | TimeInTarget_T003_Percentage_DistinguishTargetRange_mg | navigateToReports clickOnDays Manual Image Verification | 1252 1272 1273 1274 1275
SDAUIRS975 | Load the dataset. Verify app displays the percentage of historical glucose points in the reporting period that are < 70, with in target range, greater than or equal to 70 mg/dL and less than the lower target., less than or equal to 240 mg/dL and greater than the upper target,greater than 240 mg/dL. in the report as per the dataset. Verify for all reporting periods with different target ranges. Verify the color of bars are displayed correctly. | TimeInTarget_T004_Percentage_DistinguishTargetRange_mmol | navigateToReports clickOnDays Manual Image Verification | 1253 1297 1298 1299 1300
SDAUIRS976 | Load the dataset. Verify app displays the percentage of historical glucose points in the reporting period that are < 70, with in target range, greater than or equal to 70 mg/dL and less than the lower target., less than or equal to 240 mg/dL and greater than the upper target,greater than 240 mg/dL. in the report as per the dataset. Verify for all reporting periods with different target ranges. Verify the color of bars are displayed correctly. | TimeInTarget_T003_Percentage_DistinguishTargetRange_mg | navigateToReports clickOnDays Manual Image Verification | 1252 1276 1277 1278 1279
SDAUIRS976 | Load the dataset. Verify app displays the percentage of historical glucose points in the reporting period that are < 70, with in target range, greater than or equal to 70 mg/dL and less than the lower target., less than or equal to 240 mg/dL and greater than the upper target,greater than 240 mg/dL. in the report as per the dataset. Verify for all reporting periods with different target ranges. Verify the color of bars are displayed correctly. | TimeInTarget_T004_Percentage_DistinguishTargetRange_mmol | navigateToReports clickOnDays Manual Image Verification | 1253 1301 1302 1303 1304
SDAUIRS977 | Verify all percentages displayed in the Time In Target report adds to 100% | TimeInTarget_T005_Total Percentage_mg | navigateToReports clickOnDays Manual Image Verification | 1254 1313 1314 1315 1316
SDAUIRS977 | Verify all percentages displayed in the Time In Target report adds to 100% | TimeInTarget_T006_Total Percentage_mmol | navigateToReports clickOnDays Manual Image Verification | 1255 1317 1318 1319 1320 1321
SDAUIRS1124 | If the lower target is equal to 70 mg/dL, the Time In Target report shall not display the percentage of historical glucose points in the reporting period that are greater than or equal to 70 mg/dL and less than the lower target. | TimeInTarget_T003_Percentage_DistinguishTargetRange_mg | navigateToReports clickOnDays Manual Image Verification | 1252 1284 1285 1286 1287
SDAUIRS1124 | If the lower target is equal to 70 mg/dL, the Time In Target report shall not display the percentage of historical glucose points in the reporting period that are greater than or equal to 70 mg/dL and less than the lower target. | TimeInTarget_T004_Percentage_DistinguishTargetRange_mmol | navigateToReports clickOnDays Manual Image Verification | 1253 1309 1310 1311 1312

## Methods - NOT Req Specific
Test Script Name | Methods - NOT Req Specific
TimeInTarget_T001_ViewReport_ReportPeriods_TargetRange | changePhoneHourTimeFormat setTheDateAndTime navigateToHomePage loadTestData setToDefault setTargetGlucoseRange navigateToLogBook navigateToReminder currentSystemTime debugDrawerClearData
TimeInTarget_T002_DataAvailable | navigateToHomePage changePhoneHourTimeFormat setTheDateAndTime loadTestData setToDefault currentSystemTime debugDrawerClearData
TimeInTarget_T003_Percentage_DistinguishTargetRange_mg | navigateToHomePage changePhoneHourTimeFormat setTheDateAndTime loadTestData setToDefault setTargetGlucoseRange currentSystemTime debugDrawerClearData
TimeInTarget_T004_Percentage_DistinguishTargetRange_mmol | navigateToHomePage changePhoneHourTimeFormat setTheDateAndTime loadTestData setToDefault setTargetGlucoseRange currentSystemTime debugDrawerClearData
TimeInTarget_T005_Total Percentage_mg | navigateToHomePage changePhoneHourTimeFormat setTheDateAndTime loadTestData setToDefault setTargetGlucoseRange currentSystemTime debugDrawerClearData
TimeInTarget_T006_Total Percentage_mmol | navigateToHomePage changePhoneHourTimeFormat setTheDateAndTime loadTestData setToDefault setTargetGlucoseRange currentSystemTime debugDrawerClearData
SensorUsage_T001_View_Display | debugDrawerClearData navigateToHomePage changePhoneHourTimeFormat setTheDateAndTime loadTestData setToDefault navigateToLogBook navigateToReminder currentSystemTime debugDrawerClearData
SensorUsage_T002_Calculation | debugDrawerClearData navigateToHomePage changePhoneHourTimeFormat setTheDateAndTime loadTestData setToDefault currentSystemTime selectingSASMode scanMockSensor