## Traceability Matrix
Req ID | Test Cases | Test Script Name | Script Method Name | DVT Protocol Test Step ID
SDASRS1171 | Navigate to each reports. Click on share icon & select any mode. Verify app sends the screenshot via the selected mode. | General_T002_DateDisplay_ScreenCapture | navigateToReports clickOnDays saveTheSharedBitmap sendEmail | 1216 1230 1217 1218 1219 1220 1221 1222 1223 1224 1225 1226 1228 1229 1231 1232 1233 1234 1235 1236 1326 1237
SDASRS299 | Navigate to each reports from dropdown. Verify i icon is displayed. Click on i icon & verify description about the report is displayed. | General_T001_insufficientData_ViewDescription | clickOnIconAndVerifyReportMessage navigateToReports clickOk | 1208
SDASRS829* | Set SAS database to corrupt state. Open the report. Verify app notifies the user of the failure and abort further use of these data for calculations or display | General_T006_CorruptSASDatabase | navigateToReports setToSASModeToCorrupt clickOnDays verifyErrorLoadingData | 1151
SDAUIRS1148* | Disable the automatic Date & Time. Verify app displays the message when brought to foreground. | General_T004_Time_NetworkTime_Setting_Check | launch verifyAutomaticDateMessage tapOnYes automaticDateAndTime verifyHomeScreen | 1149
SDAUIRS1152 | Load the dataset multiple glucose data < 70 mg/dL & also with valid range data Navigate to each reports.select 7 days/14 days /30 days/90 days if applicable. Verify app distinguishes the sections of those traces that are < 70 mg/dL. | General_T003_Data_Distinguish | navigateToReports clickArrowOnTopLeftOfTheGraph Manual Image Verification | 1147 1238
SDAUIRS1257 | Verify each report screen displays number of days of data. Verify number of days in the reporting period shall distinguish between when those numbers are the same and when they are different. | General_T007_DistinguishDataAvailable | navigateToReports clickOnDays verifyDataAvailableColor | 1324
SDAUIRS953 | Navigate to each reports from dropdown. Verify i icon is displayed. Click on i icon & verify description about the report is displayed. | General_T001_insufficientData_ViewDescription | clickOnIconAndVerifyReportMessage navigateToReports clickOk | 1208
SDAUIRS954 | Navigate to each reports. Click on share icon & select any mode. Verify app sends the screenshot via the selected mode. Note: Verify the above step for all reports. Also select 7 days/14 days /30 days/90 days if applicable & verify the above step. | General_T002_DateDisplay_ScreenCapture | navigateToReports clickOnDays saveTheSharedBitmap sendEmail | 1216 1230 1217 1218 1219 1220 1221 1222 1223 1224 1225 1226 1228 1229 1231 1232 1233 1234 1235 1236 1326 1237
SDAUIRS955 | Navigate to each reports. Verify date/date range is displayed. Change the phone's date & select 7 days/14 days /30 days/90 day except Daily Graph and Estimated A1C. Verify the date/date range is displayed correctly. | General_T002_DateDisplay_ScreenCapture | navigateToReports clickOnDays verifyGeneralReportForAverageGlucose verifyGeneralReportForSensorUsage verifyGeneralReportForTimeInTarget verifyGeneralReportForLowGlucoseEvents verifyGeneralReportForDailyPatterns verifyGeneralReportForEstimatedA1c verifyGeneralReportForDailyGraph clickArrowOnTopLeftOfTheGraph | 1216
SDAUIRS957 | Load the dataset which has insufficient glucose data. Navigate to each reports.select 7 days/14 days /30 days/90 days if applicable. Verify app displays 'Not Enough Sensor Data'. Note: Verify the above step for all reports. | General_T001_insufficientData_ViewDescription | navigateToReports clickOnDays verifyingForNotEnoughSensorDataMessage | 1208
SDAUIRS958 | Load the dataset Change the smart device time & timezone. Verify reports displays glucose values according to their historical user time. | General_T005_Data_HistoricUserTime_TimeZoneChange | navigateToReports clickArrowOnTopLeftOfTheGraph verifyHistoricUserTimeReportFor4hrTo5dDatafile verifyHistoricUserTimeReportFor4hrTo5dAnchorageLocalPlus1hrDatafile Manual Image verification | 1150 1239 1240 1241 1242 1243 1245 1246 1247 1248 1249

## Methods - NOT Req Specific
Test Script Name | Methods - NOT Req Specific
General_T001_insufficientData_ViewDescription | setToDefault loadTestData setTheDateAndTime changePhoneHourTimeFormat setTheDate currentSystemTime debugDrawerClearData clickArrowOnTopLeftOfTheGraph
General_T002_DateDisplay_ScreenCapture | setToDefault navigateToHomePage loadTestData setTheDateAndTime changePhoneHourTimeFormat debugDrawerClearData currentSystemTime
General_T003_Data_Distinguish | changePhoneHourTimeFormat setTheDateAndTime setTargetGlucoseRange loadTestData setToDefault currentSystemTime debugDrawerClearData
General_T004_Time_NetworkTime_Setting_Check | appNetworkTimeCheck closeRunningApplication
General_T005_Data_HistoricUserTime_TimeZoneChange | changePhoneHourTimeFormat setTheDateAndTime setTargetGlucoseRange navigateToHomePage changeTimeZone launchAppWithdata appNetworkTimeCheck automaticDateAndTimeZone automaticDateAndTime loadTestData setToDefault currentSystemTime debugDrawerClearData launch
General_T006_CorruptSASDatabase | setTheDate setToDefault setSASDBToDefault currentSystemTime debugDrawerClearData
General_T007_DistinguishDataAvailable | changePhoneHourTimeFormat setTheDateAndTime loadTestData setToDefault currentSystemTime debugDrawerClearData