## Traceability Matrix
Req ID | Test Cases | Test Script Name | Script Method Name | DVT Protocol Test Step ID
SDAUIRS1175 | Navigate to About screen. Verify app displays system information in the About Screen. | SRS_T004_UOM_OS_version_NFC | navigateToAbout verifyOSVersion closeRunningApplication disableNFC verifyNFCError enableNFC launch verifyNoNFCError | 1238
SDASRS122 | Scan the sensor with below “Error Condition” selected as separate scans. “TRANSMISSION_ERROR” “RESPONSE_CORRUPT” “TEMP_HIGH” “TEMP_LOW” “TEMPORARY_PROBLEM” “REMOVED” “NOT_COMPATIBLE” Verify that the events recorded in Event Log. Verify that the Event Log displayed the events last in first order. Reboot the smart phone & verify all the recorded events exist in Event Log | Data_Checks_T002_Display_Calculation39_501 | loadTestData verifyReadyToScanPage navigateToScanResultScreen verifyMyGlucoseScreenValue navigateToLogBook clickCalanderLeftButton verifyGlucoseValueType clickBloodDropIcon clickTimepickerDone setBGR verifyBGRValueType clickAddedNote verifyLogBookDetailsScreenValue navigateToReports verifyTimeInTargetValueType verifyAverageGlucoseValueType | 1274
SDASRS307 | The software shall store Smart Device-reported timezone-change events and SAS-reported time-change events. For the second half, SAS detects and stores changes of UTC: SAS_SRS_422, SAS shall report a time change in the real-time and historic glucose data when it detects a discrepancy between the amount of time elapsed on the patch (which is monotonic) and on the phone (which is not monotonic in the presence of time changes). | UIScanning_T008_Display_TimeZoneChange_35minBtwnHG&RealT_8hrHG | verifyTimeChangeIconOnLogbookPage verifyTimeChangeIconOnLogbookDetailPage verifyTimeChangeIconOnDailyGraphPage clickOnHG450IconOnLogbookPage verifyTimeChangeIconOnHomePageForHG450 verifyIimezoneIconsForHG450 | 1176
SDASRS322 | Verify app displays the glucose results with 1 mg/dL resolution correctly as per the dataset in Home screen, Scan result screen, Logbook screen, Logbook detail screen & all applicable reports. | Data_Checks_T002_Display_Calculation39_501 | verifyReadyToScanPage navigateToScanResultScreen verifyMyGlucoseScreenValue navigateToLogBook clickCalanderLeftButton verifyGlucoseValueType clickBloodDropIcon clickTimepickerDone setBGR verifyBGRValueType clickAddedNote verifyLogBookDetailsScreenValue navigateToReports verifyTimeInTargetValueType verifyAverageGlucoseValueType | 1274
SDASRS323 | Verify app displays the glucose results with 0.1 mmol/L resolution correctly as per the dataset in Home screen, Scan result screen, Logbook screen, Logbook detail screen & all applicable reports. | Data_Checks_T002_Display_Calculation39_501 | verifyReadyToScanPage navigateToScanResultScreen verifyMyGlucoseScreenValue navigateToLogBook clickCalanderLeftButton verifyGlucoseValueType clickAddedNote verifyLogBookDetailsScreenValue navigateToReports verifyTimeInTargetValueType verifyAverageGlucoseValueType | 1274
SDASRS354* | Verify app displays the glucose values in units of measure(mg/dl ,mmol/l) , RA-insulin values(units), LA-insulin(units), Carbohydrates values(grams/servings) in units of measure in logbook screen, edit note screen, add note screen, logbook detail screen, scan result screen, home screen, settings screen & all applicable reports. Verify app displays the Sensor life in unit of measure in Home screen(DAYS,HOURS,MINUTES) | SRS_T004_UOM_OS_version_NFC | navigateToLogBook clickAddNote setLogBookTime selectAndVerifyFoodItemFromMenu setAndGetCarbAmount selectCarbohydrateUnit setAndVerifyInsulin selectExerciseAttribute clickExerciseIntensitySpinner setAndGetIntensity verifyExcerciseIntensityTimeDurationRanges swipeUpAddNoteScreen verifyCarbUnitDisplay verifyUOMNotePage clickDoneButton clickAddedNote verifyMultipleNoteLogBookDetailPage loadTestData verifyAverage navigateToReports clickOnDays Manual Image Verification verifyGlucoseUnits navigateToAppSetting clickTargetGlucoseRange clickAndCarbohydrateUnitPage selectCarbohydrateUnit navigateToHomePage navigateToScanResultScreen clickScanAddNoteButton verifyMultipleNoteDetailPage selectingSASMode editConfigurationTime verifySensorlife | 1238 1272 1273
SDAUIRS1127 | Scan the sensor with Low glucose condition & Verify app displays the notification on the result screen. Click the notification & verify app provides the option to set a time reminder to scan again in 15 mins or 30 mins. Set a time reminder to remind in 15 mins. Verify app displays the Check Glucose & running time in Reminders screen. Verify Next reminder displays Check glucose if it’s the next reminder to remind. Verify app reminds to scan the sensor in 15 mins. Set a time reminder to remind in 30mins. Verify app displays the Check Glucose & running time in Reminders screen. Verify Next reminder displays Check glucose if it’s the next reminder to remind. Verify app reminds to scan the sensor in 30 mins. Verify app allows to modify the Check Glucose time. Verify app reminds the Check glucose at the modified time. Move the app to background. Verify app reminds the Check glucose as notification at the set time | UIScanning_T001_Low_and_High_Glucose_Alarm_condition | verifyMyGlucosePage clickOnAlertIcon clickSet verifySet verifyAlarmTime selectDefaultTime setAlarmTime displaySensorTimer verifyTimerExpired verifyTimerExpiredFromBackGround | 1208 1209
SDAUIRS1128 | Scan the sensor with high glucose condition & Verify app displays the notification on the result screen. Click the notification &verify app provides the option to set a time reminder to scan again in 1 hour, 2 hours, 3 hours or 4 hours. Set a time reminder to remind in 1 hour. Verify app displays the Check Glucose & running time in Reminders screen. Verify Next reminder displays Check glucose if it’s the next reminder to remind. Verify app reminds to scan the sensor in 1 hour Similarly verify for 2 hours, 3 hours or 4 hours. Verify app allows to modify the Check Glucose time. Verify app reminds the Check glucose at the modified time. Move the app to background. Verify app reminds the Check glucose as notification at the set time. | UIScanning_T001_Low_and_High_Glucose_Alarm_condition | verifyMyGlucosePage clickOnAlertIcon clickSet verifySet verifyAlarmTime setAlarmTime displaySensorTimer verifyTimerExpired verifyTimerExpiredFromBackGround | 1208 1210
SDAUIRS1129 | Scan the sensor with projected low glucose condition & Verify app displays the notification on the result screen. Click the notification & verify app provides the option to set a time reminder to scan again in 15 mins or 30 mins. Set a time reminder to remind in 15 mins. Verify app displays the Check Glucose & running time in Reminders screen. Verify Next reminder displays Check glucose if it’s the next reminder to remind. Verify app reminds to scan the sensor in 15 mins. Set a time reminder to remind in 30mins. Verify app displays the Check Glucose & running time in Reminders screen. Verify Next reminder displays Check glucose if it’s the next reminder to remind. Verify app reminds to scan the sensor in 30 mins. | UIScanning_T002_Low_and_High_Projected_Glucose_Alarm_condition | verifyMyGlucosePage clickOnAlertIcon clickSet verifySet verifyAlarmTime displaySensorTimer setAlarmTime verifyTimerExpired | 1147 1211
SDAUIRS1130 | Scan the sensor with projected high glucose condition & Verify app displays the notification on the result screen. Click the notification & verify app provides the option to set a time reminder to scan again in 1 hour, 2 hours, 3 hours or 4 hours. Set a time reminder to remind in 1 hour. Verify app displays the Check Glucose & running time in Reminders screen. Verify Next reminder displays Check glucose if it’s the next reminder to remind. Verify app reminds to scan the sensor in 1 hour Similarly verify for 2 hours, 3 hours or 4 hours. | UIScanning_T002_Low_and_High_Projected_Glucose_Alarm_condition | verifyMyGlucosePage clickOnAlertIcon clickSet verifySet verifyAlarmTime displaySensorTimer setAlarmTime verifyTimerExpired | 1147 1212
SDAUIRS1131 | Add different real time values < 70, Scan the Mock1 sensor. Verify app distinguish a scan results < 70 mg/dL. | UIScanning_T003_Glucose_Scan_distinguish | verifyScanGlucoseValue | 1148
SDAUIRS1132 | Add different real time values > 240 mg/dL. Scan the Mock1 sensor. Verify app distinguish a scan results. | UIScanning_T003_Glucose_Scan_distinguish UIScanning_T005_MyGlucose_HG20min_BGbtw180and240_above_240 | verifyScanGlucoseValue | 1148 1150
SDAUIRS1133* | Add different real time values within target range. Scan the Mock1 sensor. Verify app distinguish a scan results that is within the target glucose range. | UIScanning_T003_Glucose_Scan_distinguish UIScanning_T005_MyGlucose_HG20min_BGbtw180and240_above_240 UIScanning_T006_BG_LessLowTarget_Greater_Or_Equal70_mg UIScanning_T007_BG_LessLowTarget_Greater_Or_Equal3.9_mmol | verifyScanGlucoseValue verifyMyGlucoseColor | 1148 1150 1151 1205
SDAUIRS1168* | Scan the sensor. Verify app displays a scan result screen containing the current Real-time glucose value, current Trend Arrow, and a glucose graph of the most recent 8 hours with timestamps converted to the user’s current timezone if the scanning is successful. Verify app displays the time of scan correctly. Verify the scan information is plotted in graph correctly. Scan the sensor. Make sure scan is unsuccessful. Verify app doesn't display a scan result screen. " Scan the sensor. Verify for all Trend arrows or no arrows. Scan the sensor. Wait for some time. Change the timezone to different timezone. Scan the sensor. Verify app displays the scan result graph as per the current timezone. Also verify the scan information is displayed correctly. | UIScanning_T004_HI_LO_MyGlucose_BG_8_hrs_NonActionableBG UIScanning_T008_Display_TimeZoneChange_35minBtwnHG&RealT_8hrHG | verifyMyGlucosePage verifyGlucosePoint verifyGlucoseGraph verifyGlucoseValueUnits navigateToMyGlucosePage | 1149 1243 1244 1245 1176 1268 1270 1271
SDAUIRS1169 | Verify app displays scan results < 40 mg/dL as “LO”. Also verify for boundary ranges 39 mg/dl, 40 mg/dl | UIScanning_T004_HI_LO_MyGlucose_BG_8_hrs_NonActionableBG | verifyHILOGlucoseValue verifyLHGlucoseGraph verifyGlucoseValue verifyLogbookColor verifyGlucosePoint verifyMyGlucosePage | 1149
SDAUIRS1170 | Verify app displays scan results > 500 mg/dL as “HI”. Also verify for boundary ranges , 500 mg/dl, 501 mg/dl | UIScanning_T004_HI_LO_MyGlucose_BG_8_hrs_NonActionableBG | verifyHILOGlucoseValue verifyLHGlucoseGraph verifyGlucoseValue verifyLogbookColor verifyGlucosePoint clickLogbookEntry verifyMyGlucosePage | 1149
SDAUIRS1176 | Add different real time values >= 70 mg/dL and < the low glucose target. Scan the Mock1 sensor. Verify app distinguish a scan results >= 70 mg/dL and < the low glucose target. Also verify boundary range at 70 mg/dL & low glucose target . | UIScanning_T006_BG_LessLowTarget_Greater_Or_Equal70_mg UIScanning_T007_BG_LessLowTarget_Greater_Or_Equal3.9_mmol | verifyScanGlucoseValue verifyGlucosePoint | 1151 1224 1262 1205
SDAUIRS1177 | Add different real time values > the high glucose target and <= 240 mg/dL.. Scan the Mock1 sensor. Verify app distinguish a scan results. Also verify for 240 mg/dL. | UIScanning_T005_MyGlucose_HG20min_BGbtw180and240_above_240 | verifyScanGlucoseValue | 1150
SDAUIRS1190 | Select Non actionable. Scan the sensor. Verify app displays the Blood & lens icon. Click on the icon & verify the displayed information & click OK. | UIScanning_T004_HI_LO_MyGlucose_BG_8_hrs_NonActionableBG | verifyNARGlucoseValue verifyNARDialog | 1149
SDAUIRS1192 | Add two or three historic values within 20 mins gap. Scan Mock_1 Sensor Verify historic glucose values are connected. Add two or three historic values > 20 mins gap. Verify historic glucose values are not connected. Add two or three historic values exactly at 20 mins gap.Verify historic glucose values are connected. | UIScanning_T005_MyGlucose_HG20min_BGbtw180and240_above_240 | verifyGlucosePoint | 1150 1218
SDAUIRS1204 | Add different real time values < 70, Scan the Mock1 sensor. Verify graph distinguish a scan results < 70 mg/dL. | UIScanning_T003_Glucose_Scan_distinguish | verifyGlucoseGraph verifyGlucosePoint | 1148 1215
SDAUIRS1205 | Add different real time values >= 70 mg/dL and < the low glucose target. Scan the Mock1 sensor. Verify graph distinguish a scan results >= 70 mg/dL and < the low glucose target. Also verify boundary range at 70 mg/dL & low glucose target . | UIScanning_T006_BG_LessLowTarget_Greater_Or_Equal70_mg UIScanning_T007_BG_LessLowTarget_Greater_Or_Equal3.9_mmol | clickLogbookEntry verifyGlucosePoint | 1151 1223 1227 1264 1205 1232 1235 1236
SDAUIRS1206 * | Add different real time values within target range. Scan the Mock1 sensor. Verify graph distinguish a scan results that is within the target glucose range. | UIScanning_T003_Glucose_Scan_distinguish UIScanning_T005_MyGlucose_HG20min_BGbtw180and240_above_240 UIScanning_T006_BG_LessLowTarget_Greater_Or_Equal70_mg UIScanning_T007_BG_LessLowTarget_Greater_Or_Equal3.9_mmol | clickLogbookEntry verifyGlucosePoint | 1148 1276 1150 1247 1151 1220 1205 1234
SDAUIRS1207 | Add different real time values > the high glucose target and <= 240 mg/dL.. Scan the Mock1 sensor. Verify graph distinguish a scan results. Also verify for 240 mg/dL. | UIScanning_T005_MyGlucose_HG20min_BGbtw180and240_above_240 | clickLogbookEntry verifyGlucosePoint | 1150 1249 1251 1253 1255
SDAUIRS1208 | Add different real time values > 240 mg/dL. Scan the Mock1 sensor. Verify graph distinguish a scan results. | UIScanning_T003_Glucose_Scan_distinguish UIScanning_T005_MyGlucose_HG20min_BGbtw180and240_above_240 | clickLogbookEntry verifyGlucosePoint | 1148 1275 1150 1257 1261 1259
SDAUIRS1209 | Add different real time values < 70, Scan the Mock1 sensor. Verify graph distinguish a scan results < 70 mg/dL. | UIScanning_T003_Glucose_Scan_distinguish | verifyGlucoseGraph verifyGlucosePoint | 1148 1214
SDAUIRS1210 | Add different real time values >= 70 mg/dL and < the low glucose target. Scan the Mock1 sensor. Verify graph distinguish a scan results >= 70 mg/dL and < the low glucose target. Also verify boundary range at 70 mg/dL & low glucose target . | UIScanning_T006_BG_LessLowTarget_Greater_Or_Equal70_mg UIScanning_T007_BG_LessLowTarget_Greater_Or_Equal3.9_mmol | verifyGlucoseGraph verifyGlucosePoint verifyMyGlucoseColor | 1151 1222 1225 1263 1205
SDAUIRS1211 * | Add different real time values within target range. Scan the Mock1 sensor. Verify graph distinguish a scan results that is within the target glucose range. | UIScanning_T003_Glucose_Scan_distinguish UIScanning_T006_BG_LessLowTarget_Greater_Or_Equal70_mg UIScanning_T007_BG_LessLowTarget_Greater_Or_Equal3.9_mmol | verifyGlucoseGraph verifyGlucosePoint verifyMyGlucoseColor | 1148 1217 1150 1246 1151 1240 1205
SDAUIRS1212 | Add different real time values > the high glucose target and <= 240 mg/dL.. Scan the Mock1 sensor. Verify graph distinguish a scan results. Also verify for 240 mg/dL. | UIScanning_T005_MyGlucose_HG20min_BGbtw180and240_above_240 | verifyGlucosePoint verifyGlucoseGraph | 1150 1248 1250 1252 1254
SDAUIRS1213 | Add different real time values > 240 mg/dL. Scan the Mock1 sensor. Verify graph distinguish a scan results. | UIScanning_T003_Glucose_Scan_distinguish UIScanning_T005_MyGlucose_HG20min_BGbtw180and240_above_240 | clickLogbookEntry verifyGlucosePoint verifyGlucoseGraph | 1148 1216 1150 1256 1260 1258
SDAUIRS1214 | Adjust the phone font size to the maximum so that scan result is not completely displayed on the phone’s screen. Verify app notifies the user. | UIScanning_T008_Display_TimeZoneChange_35minBtwnHG&RealT_8hrHG | verifyErrorDispayingResult | 1176
SDAUIRS1222* | Load the Dataset. Verify app displays the glucose results with 0.1 mmol/L resolution correctly as per the dataset in Home screen, Scan result screen, Logbook screen, Logbook detail screen & all applicable reports. | Data_Checks_T002_Display_Calculation39_501 | loadTestData verifyReadyToScanPage navigateToScanResultScreen verifyMyGlucoseScreenValue navigateToLogBook clickCalanderLeftButton verifyGlucoseValueType clickAddedNote verifyLogBookDetailsScreenValue navigateToReports verifyTimeInTargetValueType verifyAverageGlucoseValueType | 1274
SDAUIRS1227* | MMOL Build only Scan the sensor. Verify app displays the numeric glucose result digits after the decimal separator in a smaller font than the digits before the decimal separator. | UIScanning_T007_BG_LessLowTarget_Greater_Or_Equal3.9_mmol | verifyGlucosePoint | 1205 1231 1233
SDAUIRS1232 | Make sure historic glucose values time gap & scan result time gap are within 35 mins. Verify app connects historic glucose values & scan result in glucose graph Make sure historic glucose values time gap & scan result time gap are exactly 35 mins. Verify app connect historic glucose values & scan result in glucose graph Make sure historic glucose values time gap & scan result time gap are > 35 mins. Verify app doesn't connect historic glucose values & scan result in glucose graph | UIScanning_T008_Display_TimeZoneChange_35minBtwnHG&RealT_8hrHG | verifyGlucoseValueUnits navigateToMyGlucosePage | 1176 1265 1267
SDAUIRS1243 | Add a historic data. Scan the sensor Change the current timezone to different timezone. Add a historic data. Scan the sensor Verify timechange icon is displayed. Also verify the timechange icon is plotted correctly in the current timezone. | UIScanning_T008_Display_TimeZoneChange_35minBtwnHG&RealT_8hrHG | verifyGlucoseValueUnits | 1176 1266 1269
SDAUIRS1244 | Click the time change icon & verify timezone change information is displayed correctly. | UIScanning_T008_Display_TimeZoneChange_35minBtwnHG&RealT_8hrHG | clickOnTimezonesAndVerifyTimezoneChangesForHG450 clickOnTimezoneAndVerifyTimezoneChangesForHG250 | 1176
SDAUIRS847* | Verify app displays the glucose results with 1 mg/dL resolution correctly as per the dataset in Home screen, Scan result screen, Logbook screen, Logbook detail screen & all applicable reports. | Data_Checks_T002_Display_Calculation39_501 | loadTestData verifyReadyToScanPage navigateToScanResultScreen verifyMyGlucoseScreenValue navigateToLogBook clickCalanderLeftButton verifyGlucoseValueType clickBloodDropIcon clickTimepickerDone setBGR verifyBGRValueType clickAddedNote verifyLogBookDetailsScreenValue navigateToReports verifyTimeInTargetValueType verifyAverageGlucoseValueType | 1274
SDAUIRS1175 |  | SRS_T004_UOM_OS_version_NFC | verifyOSVersion | 1238

## Methods - NOT Req Specific
Test Script Name | Methods - NOT Req Specific
UIScanning_T001_Low_and_High_Glucose_Alarm_condition | debugDrawerClearData clickAlertCancel clickAlertAlarmCancel navigateToMyGlucosePage clearNotification currentSystemTime addScanData scanMockSensor closeDebugDrawer clearNotification cleanUp advanceDeviceTime selectingSASMode openPhoneSettings openBackgroundApp
UIScanning_T002_Low_and_High_Projected_Glucose_Alarm_condition | selectingSASMode clickAlertCancel debugDrawerClearData navigateToMyGlucosePage selectDefaultTime currentSystemTime addScanData scanMockSensor advanceDeviceTime clickAlertAlarmCancel clearNotification cleanUp
UIScanning_T003_Glucose_Scan_distinguish | cleanUp clickBackfromMyGlucoseScreen debugDrawerClearData addScanData scanMockSensor selectingSASMode navigateToLogBook setTargetGlucoseRange clickBackArrowORMenu
UIScanning_T004_HI_LO_MyGlucose_BG_8_hrs_NonActionableBG | setTheTime setTargetGlucoseRange selectingSASMode addScanData scanMockSensor clickBackfromMyGlucoseScreen navigateToLogBook debugDrawerClearData scanMockSensor
UIScanning_T005_MyGlucose_HG20min_BGbtw180and240_above_240 | selectingSASMode clickBackArrowORMenu navigateToHomePage addScanData scanMockSensor clickBackfromMyGlucoseScreen navigateToLogBook setTheTime loadTestData clickOnReadyToScanBar cleanUp changePhoneHourTimeFormat setTargetGlucoseRange debugDrawerClearData
UIScanning_T006_BG_LessLowTarget_Greater_Or_Equal70_mg | changePhoneHourTimeFormat debugDrawerClearData selectingSASMode setTargetGlucoseRange addScanData scanMockSensor clickBackfromMyGlucoseScreen navigateToLogBook cleanUp
UIScanning_T007_BG_LessLowTarget_Greater_Or_Equal3.9_mmol | navigateToLogBook clickBackfromMyGlucoseScreen navigateToHomePage cleanUp clickOnReadyToScanBar clickBackfromMyGlucoseScreen changePhoneHourTimeFormat setTheDateAndTime loadTestData setTargetGlucoseRange clickOnScanVlaueInLogbookPage
UIScanning_T008_Display_TimeZoneChange_35minBtwnHG&RealT_8hrHG | closeDebugDrawer navigateToHomePage navigateToLogBook openDebugDrawer clickBackArrowORMenu settingPhoneFontSize debugDrawerClearData selectingSASMode scanMockSensor addScanData timezoneCleanUp changeTimeZone setTheTime launchAppWithdata navigateToReports addMockSensorDatawithStaticTimeAndDate appNetworkTimeCheck changePhoneHourTimeFormat selectingSASMode timezoneCleanUp closeNotification
SRS_T004_UOM_OS_version_NFC | changePhoneHourTimeFormat setAndGetCarbUnit clickBackArrowORMenu deleteNoteIfAvailable doSwipeDown debugDrawerClearData closeDebugDrawer setToDefault
Data_Checks_T002_Display_Calculation39_501 | changePhoneHourTimeFormat setTheTime clickBackArrowORMenu setToDefault